## SBE MicroCATs
Model | Max. depth (m) | SMHI Nr | Part number | Serial number | Purchase Date | Latest Calibration date | Firmware version | Sensors | DO-sensor attached | Clamp size | Status | Box | Place | Station | Label on instrument | Antifouling | Cleaned | Data Downloaded | Batteries | Diagnostics done | Programmed | Ready for deployment | Comments | Date/Signature
IMP
SBE 37 IMP | 100 |  | 3747474B | 37IMP47474-6536 |  | 2012-11-08 | 2.3b | C, T, P |  |  | Stored | Lovis (2) | Lager Hylla 7B |  |  |  | Yes |  |  | Yes | No | No | ID:02; Use SeaTerm v1.59 to communicate; Slow communication; Will not respond on calls; | 2018-01-10/JKRO
SBE 37 IMP | 100 |  | 3747474B | 37IMP47474-6537 |  | 2013-03-12 |  | C, T, P |  |  | Stored | Lovis (2) | Lager Hylla 7B |  |  |  | No |  |  | No | No | No |  | 2018-08-28/MN
SBE 37 IMP | 250 |  | 37IMP.10000 | 37-10384 | 2013-02-22 | 2019-04-30 | 4.1 | C, T |  | 7/16" | Calibration | Deneb (5) | Lager Hylla 8D |  | 01 5m | No | Yes | Yes | No | Yes | No | No | Sent for calibration 2021-03-08 | 2021-03-20/MN
SBE 37 IMP | 250 |  | 37IMP.10000 | 37-10385 | 2013-02-22 | 2019-05-09 | 4.1 | C, T |  | 7/16" | Calibration | Deneb (5) | Lager Hylla 8D |  | 02 10m | No | Yes | Yes | No | Yes | No | No | Sent for calibration 2021-03-08 | 2021-03-20/MN
SBE 37 IMP | 250 |  | 37IMP.10000 | 37-10386 | 2013-02-22 | 2019-04-30 | 4.1 | C, T |  | 7/16" | Calibration | Deneb (5) | Lager Hylla 8D |  | 03 15m | No | Yes | Yes | No | Yes | No | No | Sent for calibration 2021-03-08 | 2021-03-20/MN
SBE 37 IMP | 250 |  | 37IMP.10000 | 37-10387 | 2013-02-22 | 2019-05-07 | 4.1 | C, T |  | 7/16" | Calibration | Deneb (5) | Lager Hylla 8D |  | 04 20m | No | Yes | Yes | No | Yes | No | No | Sent for calibration 2021-03-08 | 2021-03-20/MN
SBE 37 IMP | 250 |  | 37IMP.10000 | 37-10388 | 2013-02-22 | 2019-05-07 | 4.1 | C, T |  | 7/16" | Calibration | Deneb (5) | Lager Hylla 8D |  | 05 25m | No | Yes | Yes | No | Yes | No | No | Sent for calibration 2021-03-08 | 2021-03-20/MN
SBE 37 IMP | 250 |  | 37IMP.10000 | 37-10389 | 2013-02-22 | 2018-11-22 | 6.1.1 | C, T |  | 7/16" | Destroyed |  |  |  |  | Yes | No | No | Yes | No | Yes | Yes | Destroyed at Huvudskär 2020-02-11 | 2020-03-09/MN
SBE 37 IMP | 250 |  | 37IMP.10000 | 37-10742 | 2013-06-06 | 2019-05-07 | 4.3 | C, T |  | 7/16" | Calibration | P-Rigel (4) | Lager Hylla 8D |  | 06 30m | No | Yes | Yes | No | Yes | No | No | Sent for calibration 2021-03-08 | 2021-03-20/MN
SBE 37 IMP | 250 |  | 37IMP.10000 | 37-10743 | 2013-06-06 | 2019-05-07 | 4.3 | C, T |  | 7/16" | Calibration | P-Rigel (4) | Lager Hylla 8D |  | 07 40m | No | Yes | Yes | No | Yes | No | No | Sent for calibration 2021-03-09 | 2021-03-20/MN
SBE 37 IMP | 250 |  | 37IMP.10000 | 37-10744 | 2013-06-06 | 2019-05-07 | 4.3 | C, T |  | 7/16" | Calibration | P-Rigel (4) | Lager Hylla 8D |  | 08 50m | No | Yes | Yes | No | Yes | No | No | Sent for calibration 2021-03-10 | 2021-03-20/MN
SBE 37 IMP | 250 |  | 37IMP.10000 | 37-10745 | 2013-06-06 | 2019-04-30 | 4.3 | C, T |  | 7/16" | Calibration | P-Rigel (4) | Lager Hylla 8D |  | 09 60m | No | Yes | Yes | No | Yes | No | No | Sent for calibration 2021-03-11 | 2021-03-20/MN
SBE 37 IMP | 250 |  | 37IMP.10000 | 37-10746 | 2013-06-06 | 2019-04-30 | 4.3 | C, T |  | 7/16" | Calibration | Trappist (4) | Lager Hylla 8D |  | 10 70m | No | Yes | Yes | No | Yes | No | No | Sent for calibration 2021-03-12 | 2021-03-20/MN
SBE 37 IMP | 250 |  | 37IMP.10000 | 37-10747 | 2013-06-06 | 2019-04-30 | 4.3 | C, T |  | 7/16" | Deployed | Trappist (4) | Lager Hylla 8D | Huvudskär | 01 5m | Copper | Yes | Yes | Yes | Yes | Yes | Yes | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
IMP-ODO
SBE 37 IMP-ODO | 250 |  | 37IMP.10002 | 37-10395 | 2013-02-22 | 2020-07-09 | 1.1.0 | C, T, DO | 630363 | 7/16" | Deployed | Bellatrix (5) | Lager Hylla 8D | Huvudskär | 02 10m | Copper | Yes | Yes | Yes | Yes | Yes | Yes | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
SBE 37 IMP-ODO | 250 |  | 37IMP.10002 | 37-10396 | 2013-02-23 | 2020-07-09 | 1.2.0 | C, T, DO | 630364 | 7/16" | Deployed | Bellatrix (5) | Lager Hylla 8D | Huvudskär | 03 15m | Copper | Yes | Yes | Yes | Yes | Yes | Yes | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
SBE 37 IMP-ODO | 250 |  | 37IMP.10002 | 37-10740 | 2013-06-06 | 2020-10-20 | 1.1.0 | C, T, DO | 630458 | 7/16" | Deployed | Borka (5) | Lager Hylla 8D | Huvudskär | 04 20m | Copper | Yes | Yes | Yes | Yes | Yes | Yes | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
SBE 37 IMP-ODO | 250 |  | 37IMP.10002 | 37-10741 | 2013-06-06 | 2020-07-09 | 1.1.0 | C, T, DO | 630459 | 7/16" | Deployed | Borka (5) | Lager Hylla 8D | Huvudskär | 05 25m | Copper | Yes | Yes | Yes | Yes | Yes | Yes | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
SBE 37 IMP-ODO | 350 | 17092 | 37IMP.13002 | 37-20072 | 2018-06-26 | 2020-07-07 | 5.0.0 | C, T, DO, P | 63-2003 | 7/16" | Deployed | Bellatrix (5) | Lager Hylla 7C | Huvudskär | 06 30m | Copper | Yes | Yes | Yes | Yes | Yes | Yes | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
SBE 37 IMP-ODO | 350 | 17093 | 37IMP.13002 | 37-20073 | 2018-06-26 | 2020-07-07 | 5.0.0 | C, T, DO, P | 63-2004 | 7/16" | Deployed | Bellatrix (5) | Lager Hylla 7C | Huvudskär | 07 40m | Copper | Yes | Yes | Yes | Yes | Yes | Yes | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
SBE 37 IMP-ODO | 350 | 17094 | 37IMP.13002 | 37-20074 | 2018-06-26 | 2020-07-07 | 5.0.0 | C, T, DO, P | 63-2005 | 7/16" | Deployed | Bellatrix (5) | Lager Hylla 7C | Huvudskär | 08 50m | Copper | Yes | Yes | Yes | Yes | Yes | Yes | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
SBE 37 IMP-ODO | 350 | 17095 | 37IMP.13002 | 37-20100 | 2018-06-26 | 2020-07-09 | 5.0.0 | C, T, DO, P | 63-1928 | 7/16" | Deployed | Borka (5) | Lager Hylla 7C | Huvudskär | 09 60m | Copper | Yes | Yes | Yes | Yes | Yes | Yes | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
IM
SBE 37 IM | 7000 |  | 90338.006 | 37IM22422-1373 |  | 2014-04-01 | 2.2c | C, T |  |  | Stored | Polaris (4) | Lager Hylla 7D |  |  | No | Yes | Yes | No | Yes | No | No | Recovered from Tångesund 5sep2017
SBE 37 IM | 7000 |  | 90338.006 | 37IM22422-1375 |  | 2014-04-04 |  | C, T |  |  | Stored | Polaris (4) | Lager Hylla 7D |  |  |  | Yes |  |  | Yes | No | No | Havsboj; ID:10; Strange communication: sometimes it answers to commands; most of the time gibberish
SBE 37 IM | 7000 |  | 90338.006 | 37IM22422-1376 |  | 2014-04-15 | 2.2c | C, T |  |  | Stored | Rigel (5) | Lager Hylla 8D |  |  |  | Yes | Yes |  | Yes | No | No | Havsboj-Recovered in Huvudskär
SBE 37 IM | 7000 |  | 90338.006 | 37IM22422-1378 |  | 2013-01-24 | 2.2c | C, T |  |  | Stored | Mattis (4) | Lager Hylla 7B |  |  | No | Yes | Yes | Old | Yes | No | No |  | 2018-08-28/MN
SBE 37 IM | 350 |  | 90338.045 | 37IM29160-1380 |  | 2014-03-31 | 2.2c | C, T, P |  |  | Stored | Rigel (5) | Lager Hylla 7D |  |  | No | Yes | Yes | No | No | No | No | On the label says 37SM. Recovered Huvudskär 2018-03-07 ASC-file could not be converted. | 2018-03-14/MN
SBE 37 IM |  |  |  | 37IM24683-1733 |  | 2009-07-08 |  |  |  |  | Missing | [12 empty] | 2018-04-20/MN
SBE 37 IM |  |  |  | 37IM24683-1734 |  | 2009-07-09 |  |  |  |  | Missing | [12 empty] | 2018-04-20/MN
SBE 37 IM |  |  |  | 37IM24683-1735 |  | 2009-07-09 |  |  |  |  | Missing | [12 empty] | 2018-04-20/MN
SBE 37 IM |  |  |  | 37IM24683-1736 |  | 2009-07-09 |  |  |  |  | Missing | [12 empty] | 2018-04-20/MN
SBE 37 IM | 7000 |  | 90338.020 | 37IM24683-1737 |  | 2014-08-28 | 2.2c |  |  |  | Missing | [11 empty] | Not deployed at Huvudskär. Don't know where it is | 2018-04-20/MN
SBE 37 IM |  |  |  | 37IM24683-1738 |  | 2009-07-09 |  |  |  |  | Missing | [12 empty] | 2018-04-20/MN
SBE 37 IM |  |  |  | 37IM24683-1739 |  | 2009-07-09 |  |  |  |  | Missing | [12 empty] | 2018-04-20/MN
SBE 37 IM | 7000 |  | 90338.020 | 37IM24683-1740 |  | 2014-04-01 |  |  |  |  | Stored | Mattis (4) | Lager Hylla 7B |  |  | No | Yes | Yes | Old | Yes | No | No |  | 2018-08-28/MN
SBE 37 IM |  |  |  | 37IM24683-1741 |  | 2009-07-07 |  |  |  |  | Missing | [12 empty] | 2018-04-20/MN
SBE 37 IM |  |  |  | 37IM24683-1742 |  | 2009-07-07 |  |  |  |  | Missing | [12 empty] | 2018-04-20/MN
SBE 37 IM |  |  |  | 37IM24683-1743 |  | 2009-07-07 |  |  |  |  | Missing | [12 empty] | 2018-04-20/MN
SBE 37 IM |  |  |  | 37IM24683-1744 |  | 2009-07-09 |  |  |  |  | Missing | [12 empty] | 2018-04-20/MN
SBE 37 IM |  |  |  | 37IM24683-1745 |  | 2009-07-09 |  |  |  |  | Missing | [12 empty] | 2018-04-20/MN
SBE 37 IM | 7000 |  | 90338.020 | 37IM24683-1746 |  | 2013-01-24 | 2.2c |  |  |  | Stored | Mattis (4) | Lager Hylla 7B |  |  | No | Yes | Yes | Old | Yes | No | No | Packning vid TBT-rör saknas | 2018-08-28/MN
SBE 37 IM | 7000 |  | 90338.020 | 37IM24683-1747 |  | 2014-04-01 | 2.2c |  |  |  | Missing |  |  |  | A09 D45 | [7 empty] | Not deployed at Huvudskär. Don't know where it is | 2018-04-20/MN
SBE 37 IM |  |  |  | 37IM24683-1748 |  | 2009-07-09 |  | C, T, P |  |  | Missing | [12 empty] | 2018-04-20/MN
SBE 37 IM |  |  |  | 37IM24683-1749 |  | 2009-07-16 |  | C, T, P |  |  | Missing | [12 empty] | 2018-04-20/MN
SBE 37 IM |  |  |  | 37IM24683-1750 |  | 2009-07-29 |  | C, T, P |  |  | Missing | [12 empty] | 2018-04-20/MN
SBE 37 IM | 350 |  | 90338.021 | 37IM24683-1751 |  | 2014-04-15 | 2.2c | C, T, P |  |  | Stored | Rigel (5) | Lager Hylla 8D |  |  | No | Yes | Yes | No | Yes | No | No | Recovered from Tångesund 5sep2017
SBE 37 IM | 100 |  | 3731319 | 37IM31319-2855 |  | 2014-04-01 | 2.2c | C, T, P |  |  | Stored | Polaris (4) | Lager Hylla 7D |  |  | No | Yes | Yes | No | Yes | No | No | Recovered from Tångesund 5sep2017
SBE 37 IM | 7000 |  | 3731319 | 37IM31319-2856 |  | 2013-01-24 | 2.2c |  |  |  | Stored | Mattis (4) | Lager Hylla 7B |  |  | No | Yes | Yes | Old | Yes | No | No |  | 2018-08-28/MN
SBE 37 IM | 250 |  | 3747474 | 37IM47474-6521 |  | 2012-07-02 | 2.3b |  |  |  | Stored | Ronja (4) | Lager Hylla 7B |  |  | No | Yes | Yes | Old | Yes | No | No |  | 2018-08-28/MN
SBE 37 IM | 250 |  | 3747474A | 37IM47474-6529 |  | 2013-03-12 | 2.3b |  |  |  | Stored | Ronja (4) | Lager Hylla 7B |  |  | No | Yes | Yes | Old | Yes | No | No |  | 2018-08-28/MN
SBE 37 IM | 250 |  | 3747474A | 37IM47474-6530 |  | 2013-03-12 | 2.3b | C, T |  |  | Stored | Polaris (4) | Lager Hylla 7D |  |  |  | Yes | Yes |  | Yes | No | No | ID:04; Probably recovered in Huvudskär April2016;
SBE 37 IM | 250 |  | 3747474A | 37IM47474-6531 | [6 empty] | Destroyed |  | Verkstad | [9 empty] | Huvudskär | 2018-08-31/MN
SBE 37 IM | 250 |  | 3747474A | 37IM47474-6532 |  | 2013-03-12 | 2.3b |  |  |  | Stored | Ronja (4) | Lager Hylla 7B |  |  | No | Yes | Yes | Old | Yes | No | No |  | 2018-08-28/MN
SBE 37 IM | 250 |  | 3747474A | 37IM47474-6533 |  | 2013-03-12 | 2.3b |  |  |  | Stored | Ronja (4) | Lager Hylla 7B |  |  | No | Yes | Yes | Old | Yes | No | No |  | 2018-08-28/MN
SBE 37 IM | 7000 |  | 90280.3 | 37IM16530-0125 |  | 2009-12-23 |  | C, T |  |  | Stored | Birk (5) | Lager Hylla 7B |  | 4 | [5 empty] | No | No |  | 2018-08-28/MN
SBE 37 IM | 7000 |  | 90280.3 | 37IM16530-0126 |  | 2009-12-12 |  | C, T |  |  | Stored | Birk (5) | Lager Hylla 7B |  | 3 | [5 empty] | No | No |  | 2018-08-28/MN
SBE 37 IM | 7000 |  | 90280.3 | 37IM16530-0127 |  | 2009-12-12 |  | C, T |  |  | Stored | Birk (5) | Lager Hylla 7B |  | 2 | [5 empty] | No | No |  | 2018-08-28/MN
SBE 37 IM | 7000 |  | 90280.3 | 37IM16530-0128 |  | 2009-12-23 |  | C, T |  |  | Stored | Birk (5) | Lager Hylla 7B |  | 5 | [5 empty] | No | No |  | 2018-08-28/MN
SBE 37 IM | 7000 |  | 90280.3 | 37IM16530-0129 |  | 2009-12-24 |  | C, T |  |  | Stored | Birk (5) | Lager Hylla 7B |  | 6 | [5 empty] | No | No |  | 2018-08-28/MN
SBE 37 IM | 58 |  | 90280.100.3 | 37IM16530-0130 |  | 2009-12-12 |  | C, T, P |  |  | Stored |  | Lager Hylla 7B |  | 7 | [5 empty] | No | No | Alone in a box | 2018-08-28/MN
SIP
SBE 37 SIP | 7000 |  |  | 37SIP37734-3635 |  | 2004-12-12 |  | C, T |  |  | Missing | [12 empty] | 2018-04-20/MN
SBE 37 SIP | 7000 |  | 3745065 | 37SIP45065-4980 |  | 2006-10-05 |  | C, T |  |  | Stored |  | Lager Hylla 7B | [9 empty] | never used? Alone in a box | 2018-08-28/MN
SMP-ODO
SBE 37 SMP-ODO | 250 |  | 37SMP.10102 | 37-9984 |  | 2012-08-23 | 1.1.0 | C, T, DO | 630256 |  | Calibration | Aldebaran (5) | Lager Hylla 7D |  |  | No | Yes | Yes | No | Yes | No | No | Sent for calibration 2021-03-10 | 2021-03-20/MN
SBE 37 SMP-ODO | 250 |  | 37SMP.10102 | 37-9985 |  |  |  | C, T, DO | 630247 |  | Destroyed |  | Verkstad |  | Koster-mitt | [7 empty] | Destroyed at Koster 01 Dec 2016;
SBE 37 SMP-ODO | 250 |  | 37SMP.10102 | 37-9997 |  | 2012-08-23 | 1.1.0 | C, T, DO | 630257 |  | Calibration | Aldebaran (5) | Lager Hylla 7D |  |  | No |  | Yes | No |  | No | No | Sent for calibration 2021-03-10 | 2021-03-20/MN
SBE 37 SMP-ODO | 250 |  | 37SMP.10102 | 37-9999 |  | 2016-12-07 |  | C, T, DO | 630261 |  | Destroyed |  | Verkstad |  | Tangesund 10m | [7 empty] | Destroyed at Tångesund 5 sep 2017
SBE 37 SMP-ODO | 250 |  | 37SMP.10102 | 37-10000 |  | 2016-01-09 | 1.1.0 | C, T, DO | 630242 |  | Calibration | Aldebaran (5) | Lager Hylla 7D |  |  | No | Yes | Yes | No |  | No | No | Sent for calibration 2021-03-10 | 2021-03-20/MN
SBE 37 SMP-ODO | 250 |  | 37SMP.10102 | 37-10001 |  | 2012-08-24 | 1.1.0 | C, T, DO | 630246 |  | Calibration | Aldebaran (5) | Lager Hylla 7D |  |  | No | Yes | Yes | No | Yes | No | No | Sent for calibration 2021-03-10 | 2021-03-20/MN
SBE 37 SMP-ODO | 250 |  | 37SMP.10102 | 37-10002 |  | 2012-08-24 | 1.1.0 | C, T, DO | 630260 |  | Calibration | Aldebaran (5) | Lager Hylla 7D |  |  | No | Yes | Yes | No |  | No | No | Sent for calibration 2021-03-10 | 2021-03-20/MN
SBE 37 SMP-ODO | 250 |  | 37SMP.10102 | 37-10017 |  | 2020-05-26 | 1.1.0 | C, T, DO | 630259 | 12mm | Deployed |  |  | Hanöbukten | Hanöbukten | Copper | Yes | Yes | Yes | Yes | Yes | Yes | Deployed at Hanöbukten 2021-03-20 | 2021-03-20/MN
SBE 37 SMP-ODO | 250 |  | 37SMP.10102 | 37-10018 |  | 2020-05-26 | 1.1.0 | C, T, DO | 630245 | 12mm | Deployed |  |  | L9 | L9 | Copper | Yes | Yes | Yes | Yes | Yes | Yes | Deployed L9 2020-12-13 | 2020-12-17/MN
SBE 37 SMP-ODO | 250 |  | 37SMP.10102 | 37-10823 |  | 2018-10-02 | 1.1.0 | C, T, DO | 630500 | No clamps | Deployed | Vega (2) | Lager Hylla 7C | P22/3892 | P22/3892 | Copper | Yes | Yes | Yes | Yes | Yes | Yes | Deployed at P22/3892 2021-03-21 | 2021-03-21/MN
SBE 37 SMP-ODO | 250 |  | 37SMP.10102 | 37-10824 |  | 2020-05-28 | 1.1.0 | C, T, DO | 630501 | 12mm | Recovered | Adela (4) | Lager Hylla 7C |  |  | Copper | Yes | Yes | No | Yes | Yes | Yes | Recovered at Hanöbukten 2021-03-20 | 2021-03-20/MN
SBE 37 SMP-ODO | 250 |  | 37SMP.10102 | 37-10825 |  | 2020-06-04 | 1.1.0 | C, T, DO | 630503 | 12mm | Deployed | Adela (4) | Lager Hylla 7C | Understen | Understen | Copper | Yes | Yes | Yes | Yes | Yes | Yes | Deployed at Understen 2020-12-05 | 2020-12-07/MN
SBE 37 SMP-ODO | 250 |  | 37SMP.10102 | 37-10826 |  | 2020-05-28 | 1.1.0 | C, T, DO | 630504 | No clamps | Stored | Adela (4) | Lager Hylla 7C |  |  | No | Yes | Yes | No | Yes | No | No | Back from calibration 2020-06-15 | 2020-06-16/MN
SBE 37 SMP-ODO | 250 |  | 37SMP.10102 | 37-10827 |  | 2020-05-28 | 1.1.0 | C, T, DO | 630505 | No clamps | Stored | Adela (4) | Lager Hylla 7C |  |  | No | Yes | Yes | No | Yes | No | No | Back from calibration 2020-06-15 | 2020-06-16/MN
SBE 37 SMP-ODO | 100 |  | 37SMP.12102 | 37-12791 |  | 2018-10-02 | 2.3.1 | C, T DO, P | 630891 | 10mm | Deployed | Alpha Centauri (5) | Lager Hylla 7C | Flinten-7 | Flinten-7 2m | Copper | Yes | Yes | Yes | Yes | Yes | Yes | Deployed Flinten-7 2020-09-01 | 2020-10-09/MN
SBE 37 SMP-ODO | 100 |  | 37SMP.12102 | 37-12797 |  | 2018-10-02 | 2.3.1 | C, T DO, P | 630896 |  | Stored | Alpha Centauri (5) | Lager Hylla 7C |  |  | No | Yes | Yes | No |  | No | No | Back from calibration 2018-11-19 | 2018-11-19/MN
SBE 37 SMP-ODO | 100 |  | 37SMP.12102 | 37-12798 |  | 2018-10-02 | 2.3.1 | C, T DO, P | 630898 |  | Stored | Alpha Centauri (5) | Lager Hylla 7C |  |  | No | Yes | Yes | No | Yes | No | No | Back from calibration 2018-11-19 | 2018-11-19/MN
SM - APV-instrument
SBE 37 SM | 100 |  |  | 37SM55059-7179 |  |  |  | C, T, P |  |  | Stored | [11 empty] | 1 clamp only. Not i Furud, maybe in another box? | 2019-08-14/MN
SBE 37 SM | 100 |  | 3755059B | 37SM55059-7180 |  | 2016-01-13 | 3.1 | C, T, P |  | No clamps | Stored | Furud (5) | Lager Hylla 7D |  |  | No | Yes | No | No | Yes | No | No | Recovered from Öregrund 2020-09-15 | 2020-09-23/RS
SBE 37 SM | 100 |  | 3755059B | 37SM55059-7181 |  | 2016-01-13 | 3.1 | C, T, P |  | No clamps | Stored | Sirius (4) | Lager Hylla 7C |  |  | No | Yes | No | No | Yes | No | No | For calibration - Labels on the instrument damaged | 2020-09-23/RS
SBE 37 SM | 100 |  | 3755059B | 37SM55059-7182 |  | 2016-02-17 | 3.1 | C, T, P |  | No clamps | Stored | Furud (5) | Lager Hylla 7D |  |  | No | Yes | No | No | Yes | No | No | Recovered from Öregrund 2020-09-15 | 2020-09-23/RS
SBE 37 SM | 100 |  | 3755059B | 37SM55059-7184 |  | 2016-01-13 | 3.1 | C, T, P |  | No clamps | Stored | Sirius (4) | Lager Hylla 7C |  |  | No | Yes | No | No | Yes | No | No | Recovered from Öregrund 2020-09-15 | 2020-09-23/RS
SBE 37 SM | 100 |  | 3755059B | 37SM55059-7185 |  | 2016-02-17 |  | C, T, P |  |  | Stored | Furud (5) | Lager Hylla 7D |  |  | No | Yes | No | No | No | No | No | Can not connect to instrument, Service? | 2019-08-15/MN
SBE 37 SM | 100 |  | 3755059B | 37SM55059-7186 |  | 2009-06-16 | 3.0h | C, T, P |  | No clamps | Stored | Sirius (5) | Lager Hylla 7D |  |  | No | Yes | No | No | Yes | No | No | Recovered from Öregrund 2020-09-15 | 2020-09-23/RS
SBE 37 SM | 100 |  | 3755059B | 37SM55059-7187 |  | 2016-01-13 | 3.1 | C, T, P |  | No clamps | Stored | Vega (2) | Lager Hylla 7C |  |  | No | Yes | Yes | No | Yes | No | No | Recovered at Flinten 2020-09-01. | 2020-09-16/MN
SBE 37 SM | 250 |  | 3755059C | 37SM55059-7188 |  | 2009-06-16 | 3.0h | C, T, P |  | No clamps | Stored | Furud (5) | Lager Hylla 7D |  |  | No | Yes | No | No | Yes | No | No | Recovered from Öregrund 2020-09-15 | 2020-09-23/RS
SBE 37 SM | 250 |  | 3755059C | 37SM55059-7191 |  | 2016-02-15 | 3.1 | C, T, P |  | No clamps | Stored | Furud (5) | Lager Hylla 7C |  |  | No | Yes | No | No | Yes | No | No | Recovered from Öregrund 2020-09-15 | 2020-09-23/RS
SBE 37 SM | 250 |  | 3755059C | 37SM55059-7193 |  | 2016-01-13 | 3.1 | C, T, P |  | No clamps | Stored | Sirius (4) | Lager Hylla 7C |  |  | No | Yes | No | No | Yes | No | No | Recovered from Öregrund 2020-09-15 | 2020-09-23/RS
SBE 37 SM | 250 |  | 3755059B | 37SM55059-7194 |  | 2016-01-13 | 3.1 | C, T, P |  | No clamps | Stored | Sirius (4) | Lager Hylla 7C |  |  | No | Yes | No | No | Yes | No | No | Recovered from Öregrund 2020-09-15 | 2020-09-23/RS
SBE 44 Underwater Inductive Modem (UIM)
SBE 44 UIM | 7000 |  | 90343.005 | 4425161-0017 |  | 2012-08-27 |  |  |  |  | Stored | Undis (5) | Lager Hylla 7B | [9 empty] | Check for calibration; | 2018-08-28/MN
SBE 44 UIM |  |  | 4433362 | 4433362-0059 | [6 empty] | Stored | Undis (5) | Lager Hylla 7B | [10 empty] | 2018-08-31/MN
SBE 44 UIM |  |  |  | 4438141-0066 |  | 2012-08-27 |  |  |  |  | Stored | Undis (5) | Lager Hylla 7B | [10 empty] | 2018-08-28/MN
SBE 52-MP Moored Profiler
SBE 52 | 100 |  |  | 5237734-0001 |  | 2012-07-31 |  | C, T DO, P | 430079 |  | Missing

## SBE SeaCATs
Model | Max. depth (m) | Part number | Serial number | Purchase Date | Latest Calibration date | Sensors | DOS Attached | Pump Attached | T sensor attached | C sensor attached | WETlabs sensor attached | WETlabs sensor attached | Status | Place | Station | Label on instrument | Antifouling | Cleaned | Data Downloaded | Batteries | Diagnostics done | Programmed | Ready for deployment | Comments | Date/Signature
16plus V2
SBE 16plus | 600 | 90641.001 | 16P57088-6475 |  | 2014-04-08 | C, T, DO | 431748 | 55482 |  |  |  |  | Stored | Lager Hylla 8B | [6 empty] | Yes |  |  | Batteries replaced (9 D-cell Duracell Procell), O-rings ok, | 2018-03-07/Jkro
SBE 16plus | 600 | 90641.001 | 16P60193-6599 |  | 2016-01-15 | C, T, DO | 431898 | 55672 |  |  |  |  | Stored | Lager Hylla 8B | [9 empty] | Batteries ok | 2018-03-07/Jkro
SBE 16plus | 600 | 90641.001 | 16P71508-7280 | 2013-02-22 | 2019-05-09 | C, T, DO | 630382 | 56902 |  |  |  |  | Calibration | Lager Hylla 8B |  |  | No | Yes |  | Yes | Yes | No | No | Sent for calibration 2021-03-08 | 2021-03-20/MN
SBE 16plus | 600 | 90641.001 | 16P74016-7331 | 2013-06-06 | 2020-07-09 | C, T, DO, Chl, Turb, PC | 630461 | 57050 |  |  | FLNTUS-1527 | FLPCS-2570 | Deployed | Lager Hylla 8B |  | ODAS03 Huvudskär | Copper | Yes |  | Yes | No | No | No | Back from service 2020-10. The 6-pin I/O bulkhead connector was found to have corrosion damage. New installed. Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
19plus V2
SBE 19plus | 350 | 90635.004 | 19P47474-6164 |  | 2012-08-08 | C, T DO, P | [6 empty] | Stored | Svea SMHI Store | [9 empty] | Comes with cage; Check before deployment
9plus "Utsjö-CTD"
SBE 9plus |  |  | 09P33999-0745 | 2004-04-13 | 2018-07-03 | C ,T, DO, P, Chl, Turb, PC, Irradiance, Depth |  | 53815 | 03P4361 | 42994 | FLNTUS-2587 |  | Stored | Svea SMHI Store | Triaxus |  | No | [6 empty] | Pressure SN0745; Altimeter: SN53447 | 200913/Jkro
[8 empty] | 54511 | 03P4387 | 42997
SBE 9plus | 1368 | 90536.067 | 09P61534-1044 |  | 2020-01-31 | C, T, DO, P |  | 56060 | 03P5351 | 43943 |  |  | Stored | Svea Scientific Store |  |  | No | [6 empty] | 04-3943 and 03-5355 mounted on 09-1387 , June 2020. | 200611/AWW
[8 empty] | 56061 | 03P5355 | 43941
SBE 9plus | 1368 | 09P.1G242 | 09-1387 | 2019-07-01 | 2019-06-10 | C, T, DO, Chl, Turb, PC | 43-3823 | 05-9782 | 03-6378 | 04-4824 | FLNTURTD-2288 | FLPCRTD-2299 | Deployed | Svea CTD Hangar | Rosett | [8 empty] | 04-4829 and 03-6384 has been changed on June expedition on Svea due to drift in the 04-4829. Anna has contacted SBE about a possible warranty claim. | 200611/AWW
[7 empty] | 43-3829 | 05-9868 | 03-6384 | 04-4829 | [14 empty] | 200611/AWW
SBE11plus Deck Unit
SBE 11plus |  |  | 11P33999-0685 | 2004-04-13
SBE 11plus |  | 90876 | 11-1158 | 2019-07-01 | [20 empty] | 2019-08-21/Jkro
SBE32 Carousel
SBE 32 | [12 empty] | Stored | Svea SMHI Store | [9 empty] | 12 flaskors, gamla typen av kontakt, finns i trälåda med kabel | 2020-01-29/Jkro
SBE 32 (ägs av SLU?) | [12 empty] | Deployed | Svea CTD Hangar | Rosett | [8 empty] | 24 flaskor, ny typ av kontakt | 2020-01-29/Jkro

## SBE Dissolved Oxygen Sensors
Model | Max. depth (m) | Part number | Serial number | Purchase Date | Latest Calibration date | Attached to MicroCAT | Attached to SeaCAT | Comments | Type of connector | Date/Signature
SBE 63 Kan bestyckas på IMP, SMP & SeaCATs; Max pressure according to Sea-Bird Pressure Test
SBE 63 ODO | 870 | 90748 | 630242 |  | 2016-01-07 | 37-10000
SBE 63 ODO |  | 90748 | 630245 |  | 2020-08-13 | 37-10018 |  |  |  | 2020-10-09/MN
SBE 63 ODO |  | 90748 | 630246 |  | 2012-09-15 | 37-10001
SBE 63 ODO |  | 90748 | 630247 |  |  | 37-9985 |  | Destroyed in Koster; needs to be send for service and calibration; to be attached to a mount and used with a SeaCAT
SBE 63 ODO |  | 90748 | 630256 |  | 2012-09-12 | 37-9984
SBE 63 ODO |  | 90748 | 630257 |  | 2012-09-13 | 37-9997
SBE 63 ODO | 870 | 90748 | 630259 |  | 2020-07-16 | 37-10017 |  |  |  | 2020-10-09/MN
SBE 63 ODO |  | 90748 | 630260 |  | 2012-09-15 | 37-10002
SBE 63 ODO |  | 90748 | 630261 |  | 2016-12-02 | 37-9999 |  | Destroyed at Tångesund 09-2017
SBE 63 ODO |  | 90748 | 630363 | 2013-02-22 | 2020-08-13 | 37-10395 |  |  |  | 2020-10-14/MN
SBE 63 ODO |  | 90748 | 630364 | 2013-02-22 | 2020-08-13 | 37-10396 |  |  |  | 2020-10-14/MN
SBE 63 ODO | 600 | 90802 | 630382 | 2013-02-22 | 2019-05-21 |  | 16P71508-7280 | Back from calibration 2019-06-27. |  | 2020-03-10/MN
SBE 63 ODO |  | 90843 | 630458 | 2013-06-06 | 2020-08-06 | 37-10740 |  |  |  | 2020-10-14/MN
SBE 63 ODO |  | 90843 | 630459 | 2013-06-06 | 2020-08-06 | 37-10741 |  |  |  | 2020-10-14/MN
SBE 63 ODO | 600 | 90840 | 630461 | 2013-06-06 | 2018-07-25 |  | 16P74016-7331 | Single point calibration, 2020-09-29 |  | 2020-10-14/MN
SBE 63 ODO |  | 90843 | 630500 |  | 2014-12-13 | 37-10823 |  | Single point calibration, 2018-10-02 |  | 2020-06-16/MN
SBE 63 ODO |  | 90843 | 630501 |  | 2015-12-31 | 37-10824 |  | Single point calibration, 2020-05-28 |  | 2020-06-16/MN
SBE 63 ODO |  | 90843 | 630503 |  | 2015-12-31 | 37-10825 |  | Single point calibration, 2020-06-04 |  | 2020-06-16/MN
SBE 63 ODO |  | 90843 | 630504 |  | 2015-12-31 | 37-10826 |  | Single point calibration, 2020-05-28 |  | 2020-06-16/MN
SBE 63 ODO |  | 90843 | 630505 |  | 2015-12-31 | 37-10827 |  | Single point calibration, 2020-05-28 |  | 2020-06-16/MN
SBE 63 ODO |  | 63.11 | 630891 |  | 2014-10-03 | 37-12791 |  | Single point calibration, 2018-10-02 |  | 2020-06-16/MN
SBE 63 ODO |  | 63.11 | 630896 |  | 2014-10-09 | 37-12797 |  | Single point calibration, 2018-10-02 |  | 2020-06-16/MN
SBE 63 ODO |  | 63.11 | 630898 |  | 2014-10-09 | 37-12798 |  | Single point calibration, 2018-10-02 |  | 2020-06-16/MN
SBE 63 ODO |  | 63.11 | 63-1928 | 2018-06-26 | 2020-08-06 | 37-20100 |  |  |  | 2020-10-14/MN
SBE 63 ODO |  | 63.11 | 63-2003 | 2018-06-26 | 2020-08-06 | 37-20072 |  |  |  | 2020-10-14/MN
SBE 63 ODO |  | 63.11 | 63-2004 | 2018-06-26 | 2020-08-06 | 37-20073 |  |  |  | 2020-10-14/MN
SBE 63 ODO |  | 63.11 | 63-2005 | 2018-06-26 | 2020-08-06 | 37-20074 |  |  |  | 2020-10-14/MN
SBE 43 Kan bestyckas på SeaCATs
SBE 43 | 600 | 90613 | 43-1748 |  | 2016-10-28 |  | 16P57088-6475 | Monterade på SBE16 för havsboj |  | 190411/AW
SBE 43 | 600 | 90613 | 43-1898 |  | 2016-01-15 |  | 16P60193-6599 | Monterade på SBE16 för havsboj |  | 190411/AW
SBE 43 | 600 | 90599 | 43-1823 |  | 2020-03-13 |  | 19P47474-6164 | Monterad på SBE19 - sept 2020. | Wet-pluggable (MCBH) | 2020-09-29/Jkro
SBE 43 | 7000 | 90572 | 43-1243 |  | 2018-06-01 |  |  | Svea dec 2019, i SBE9-lådan | Wet-pluggable (MCBH) | 2019-11-26/Jkro
SBE 43 | 7000 | 90546 | 43-0732 |  | 2019-04-26 |  |  | SVEA Nov 2019 | Wet-pluggable (MCBH) | 2019-11-26/Jkro
SBE 43 | 7000 | 90599.2 | 43-2126 |  | 2017-04-14 |  |  | Svea dec 2019 i SBE9-lådan | Wet-pluggable (MCBH) | 2019-11-26/Jkro
SBE 43 | 7000 | 90599.2 | 43-2127 |  | 2019-06-14 |  |  | SVEA Nov 2019 | Wet-pluggable (MCBH) | 2019-11-26/Jkro
SBE 43F |  |  | 43-0079 |  | 2012-08-14 | 5237734-0001
SBE 43 | 7000 | 43.2211 | 43-3823 | 2019-07-01 | 2019-06-06 |  | 09-1387 | Label on cable: OXY 1 | Wet-pluggable (MCBH) | 2019-08-21/Jkro
SBE 43 | 7000 | 43.2211 | 43-3829 | 2019-07-01 | 2019-06-06 |  | 09-1387 | Label on cable: OXY 2 | Wet-pluggable (MCBH) | 2019-08-21/Jkro

## SBE Pumps
Model | Max. depth (m) | Part number | Serial number | Pump speed, (RPM) | Type of connector | Purchase Date | Latest Calibration date | Label on instrument | Attached to SeaCAT | Comments | Date/Signature
SBE 05
SBE 05T | 10500 | 90160 | 05-3855 | 3K | XSG - i.e. old type |  | 2020-01-16 |  |  | At the office | 2019-11-26/Jkro
SBE 05 |  |  | 55672 |  |  |  | 2016-02-25 |  | 16P60193-6599
SBE 05T | 10500 | 90736 | 05-6801 | 2K |  |  | 2012-10-23 |  | 19P47474-6164
SBE 05M |  |  | 05M1108 |  |  |  | 2012-10-31
SBE 05P | 600 | 90617 | 55482 |  |  |  | 2014-05-19 |  | 16P57088-6475
SBE 05P | 600 | 90617 | 56902 |  |  | 2013-02-22 | 2019-05-09 |  | 16P71508-7280 |  | 2019-05-21/MN
SBE 05P | 600 | 90617 | 57050 |  |  | 2013-06-06 | 2020-09-29 |  | 16P74016-7331 |  | 2020-10-14/MN
SBE 05T |  |  | 53815 |  |  |  | 2018-05-14 |  | 09P33999-0745
SBE 05T |  |  | 54511 |  |  |  | 2018-05-14 |  | 09P33999-0745
SBE 05T | 10500 | 90544 | 56060 |  |  |  | 2016-12-15 | Primary pump | 09P61534-1044
SBE 05T | 10500 | 90544 | 56061 |  |  |  | 2016-12-15 | Secondary pump | 09P61534-1044
SBE 05P |  |  | 55881 |  |  |  | 2011-01-06
SBE 05P |  |  | 55882 |  |  |  | 2011-01-06
SBE 05T | 10500 | 05T.22130 | 05-9782 | 3K | MCBH | 2019-07-01 | 05-12-2018 |  | 09-1387 | Label on cable: SBE 5 | 2019-08-21/Jkro
SBE 05T | 10500 | 05T.22130 | 05-9868 | 3K | MCBH | 2019-07-01 | 01-04-2019 |  | 09-1387 | Label on cable: SBE 5 | 2019-08-21/Jkro
|  |  |  | link to connectors: | https://www.seabird.com/connector-types

## SBE T&C-sensors
Model | Max. depth (m) | Part number | Serial number | Purchase Date | Latest Calibration date | Label on instrument | Attached to SeaCAT | Comments | Date/Signature
SBE 3plus (Temperature Sensor) Used with SBE 9plus SeaCATs
SBE 3plus | 6800 | 03P.220 | 03-6378 | 2019-07-01 | 2019-05-08 |  | 09-1387 | Label on cable: SBE 3 | 2019-08-21/Jkro
SBE 3plus | 6800 | 03P.220 | 03-6384 | 2019-07-01 | 2019-05-04 |  |  | Label on cable: SECONDARY. Pay kontoret 2020-09-29. | 2020-09-29/Jkro
SBE 3plus |  |  | 03P4361 |  | 2020-02-06 | Primary | 09P33999-0745
SBE 3plus |  |  | 03P4387 |  | 2020-01-15 | Secondary | 09P33999-0745
SBE 3plus | 6800 | 90464 | 03P5351 |  | 2020-02-06 | Primary | 09P61534-1044 | även kalibrerad hos RISE 2020-03-24 | 2020-09-29/Jkro
SBE 3plus | 6800 | 90464 | 03P5355 |  | 2020-02-06 | Secondary | 09-1387 | även kalibrerad hos RISE 2020-03-31, ersatte 6384 hos 1387 juni 2020 pga problem. | 2020-09-29/Jkro
SBE 35 |  |  | 3540581-0054 |  | 2004-05-17
SBE 4C (Conductivity Sensor) Used with SBE 9plus SeaCATs
SBE 4C | 6800 | 04C.220 | 04-4824 | 2019-07-01 | 2019-05-08 |  | 09-1387 | lable on cable: SBE 4 | 2019-08-21/Jkro
SBE 4C | 6800 | 04C.220 | 04-4829 | 2019-07-01 | 2019-05-08 |  |  | lable on cable: SECONDARY. På kontoret 2020-09-29. | 2019-08-21/Jkro
SBE 4C |  |  | 42994 |  | 2020-01-14 | Primary | 09P33999-0745
SBE 4C |  |  | 42997 |  | 2020-01-14 | Secondary | 09P33999-0745
SBE 4C | 6800 | 90467 | 43941 |  | 2020-01-23 | Primary | 09P61534-1044
SBE 4C | 6800 | 90467 | 43943 |  | 2020-01-23 | Secondary | 09-1387 | ersatte 4829 hos 1387 juni 2020 pga. problem. | 2020-09-29/Jkro
SBE 38 (Digital Thermometer)
SBE 38 | 10500 |  | 3860388-0569 |  | 2010-07-13 |  |  | oklar placering
SBE 38 | 10500 | 90299.6 | 3867918-0656 |  | 2020-03-13 |  |  | Sensor for Tavastland. Calibrated at IOW in spring 2020. Placed on Tavastland before change of instruments.KONTROLLERA SÅ KOEFFICIENTERNA ÄR INLAGDA I SENSORN!!!! | 2020-11-09/AWW
SBE 38 | 10500 |  | 3854077-0507 |  | 2009-07-21 |  |  | oklar placering
SBE 38 | 10500 | 90299.6 | 3852713-0451 |  | 2018-10-16 |  |  | Sensor for Tavastland. Installed at Tavastland 20-01-10 | 200206/AWW

## SBE Thermosalinograph
Model | Max. depth (m) | Part number | Serial number | Latest Calibration date | Sensors | Comments | Date/Signature
SBE 45 Thermosalinograph
SBE 45 |  |  | 4525478-0027 | 2016-01-21 | C, T | Sensor for Tavastland. Installed at Tavastland 20-01-10. | 200117/Awwr
SBE 45 |  |  | 4524843-0020 | 2012-11-14 | C, T
SBE 45 |  |  | 4554077-0268 | 2009-07-23 | C, T | Förrådet: Pall TransPaper: In smal black pelicase
SBE 46 | 50 psi | 4567918 | 4567918-0409 | 2020-03-13 | C, T | Sensor for Tavastland. Calibrated at IOW in spring 2020. Placed on Tavastland before change of instruments. KONTROLLERA SÅ KOEFFICIENTERNA ÄR INLAGDA I SENSORN!!!! | 20-11-09/Awwr

## WETLabs
Model | Max. depth (m) | SMHI Nr | Serial number | Purchase Date | Latest Calibration date | Sensors | Status | Box | Place | Station | Attached to SeaCAT | Mooring type | Wiper | Additional Information | Comments | Date/Signature
WETLabs Fluorometer
WETLabs | 300m |  | FLNTUS-660 |  | 2014-04-23 | Chl, Turb | Destroyed |  |  |  |  | Havsboj | Yes | Chlorophyll (470/695nm); Turbidity (700nm); FLNTUS—The capabilities of the FLNTU with an integrated anti-fouling Bio-wiper™. | Sent for Calibration and doomed by Wetlabs. No longer with us, RIP! | 2019-03-11/AWW
WETLabs | 300m |  | FLNTUS-1527 |  | 2019-03-04 | Chl, Turb | Deployed | WL3 (3) | Lager Hylla 7D |  | 16P74016-7331 | Havsboj | Yes | Chlorophyll (470/695nm); Turbidity (700nm); FLNTUS—The capabilities of the FLNTU with an integrated anti-fouling Bio-wiper™; | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
WETLabs | 300m |  | FLS-012 |  | 2012-10-31 | Chl | [6 empty] | Yes | Chlorophyll(Chl) (470/695nm); FLS—Provides the capabilities of the FL with an integrated anti-fouling Bio-wiper™.
WETLabs | 300m |  | FLNTUS-1100 |  | 2019-06-13/Feb2020 SYKE | Chl, Turb | Stored | WL3 (3) | Lager Hylla 7D |  |  | Havsboj | Yes | Chlorophyll (470/695nm); Turbidity (700nm); FLNTUS—The capabilities of the FLNTU with an integrated anti-fouling Bio-wiper™ | Recovered Huvudskär 2020-12-10 | 2020-12-17/MN
WETLabs | 300m |  | FLNTUS-1098 |  | 21-Apr-2014/01-Feb-2017 | Chl, Turb | Stored | WL3 (3) |  |  |  | Havsboj | Yes | Chlorophyll (470/695nm); Turbidity (700nm); FLNTUS—The capabilities of the FLNTU with an integrated anti-fouling Bio-wiper™ |  | 2019-04-26/MN
WETLabs | 300m |  | FLNTUS-277 |  | 2012-07-11 | Chl, Turb | Stored | WL1 (3) | Lager Hylla 7D |  |  | Havsboj | Yes | Chlorophyll (470/695nm); Turbidity (700nm); FLNTUS—The capabilities of the FLNTU with an integrated anti-fouling Bio-wiper™; | Needs a check; | 2018-08-31/MN
WETLabs | 300m |  | FLNTUS-1815 |  | 2010-04-22 | Chl, Turb | Stored | WL1 (3) | Lager Hylla 7D |  |  | Havsboj | Yes | Chlorophyll (470/695nm); Turbidity (700nm); FLNTUS—The capabilities of the FLNTU with an integrated anti-fouling Bio-wiper™ | Needs a check and a proper cleaning | 2018-08-21/MN
WETLabs | 300m |  | FLPCS-2570 |  | 2020-08-28 | PC | Deployed | WL1 (3) |  |  | 16P74016-7331 | Havsboj | Yes | Phycocyanin (PC) (630/680nm); FLS—Provides the capabilities of the FL with an integrated anti-fouling Bio-wiper™. | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
WETLabs | 6000m |  | FLPCRTD-2299 |  | 2011-06-24 | PC | Deployed |  | Svea CTD Hangar | Rosett | 09-1387 | Rosett utsjöövervakningen | No | Phycocyanin (PC) (630/680nm); FL(RT)—Provides analog or RS-232 serial output with 16,300-count (approximate) range. “Real Time” instruments provide continuous operation when powered; FL(RT)D—Provides the capabilities of the FL(RT) with 6,000-meter depth rating. | Används inom miljöövervakningen | 2019-03-14/AWW
WETLabs | 6000m |  | FLPCRTD-2290 |  | 2014-08-13 | PC | Stored |  |  |  |  | Rosett utsjöövervakningen | No | Phycocyanin (PC) (630/680nm); FL(RT)—Provides analog or RS-232 serial output with 16,300-count (approximate) range. “Real Time” instruments provide continuous operation when powered; FL(RT)D—Provides the capabilities of the FL(RT) with 6,000-meter depth rating. | needs a checkup, 170707/DBSJ. At Macartney and mounted on Triaxus | 2019-03-14/AWW
WETLabs | 600m | 15405 | FLNTURT-1899 |  | Feb_2020_SYKE | Chl, Turb | Deployed |  | Svea SMHI Store |  |  | Rosett utsjöövervakningen | No | Chlorophyll (470/695nm); Turbidity (700nm); FLNTU(RT)— Analog and RS-232 serial output with 4,000-count range. Unit operates continuously when power is supplied. | To be sent to Svea June 2020/MN | 2020-05-15/MN
WETLabs | 600m |  | FLNTURT-115 | 2004-04-13 | 2013-05-22/Feb 2018_SMHI | Chl, Turb | Deployed |  |  |  | 09-1387 | Rosett utsjöövervakningen | No | Chlorophyll (470/695nm); Turbidity (700nm); FLNTU(RT)— Analog and RS-232 serial output with 4,000-count range. Unit operates continuously when power is supplied. | Undergoes checking before deployment; | 200221/AWWr
WETLabs | 6000m |  | FLNTURTD-2288 |  | 2014-08-19/Feb 2018_SMHI/Feb_2020_SYKE | Chl, Turb | Stored | WL2 (3) | Svea CTD Hangar | Rosett |  | Rosett utsjöövervakningen | No | Chlorophyll (470/695nm); Turbidity (700nm); FLNTU(RT)— Analog and RS-232 serial output with 4,000-count range. Unit operates continuously when power is supplied. |  | 200221/AWWr
WETLabs | 300m |  | FLNTUS-2587 |  | 2017-01-01/Feb_2020_SYKE | Chl, Turb | Deployed |  | Tavastland | Ferrybox |  | Ferrybox | No | Chlorophyll (470/695nm); Turbidity (700nm); FLNTU(RT)— Analog and RS-232 serial output with 4,000-count range. Unit operates continuously when power is supplied. |  | 200221/AWWr
WETLabs | 600 m |  | FLNTURT-5359 |  | 2018-11-11 | Chl, Turb | Deployed |  | Svea Sea Water lab | Ferrybox |  | Ferrybox | No | Chlorophyll (470/695nm); Turbidity (700nm); FLNTU(RT)— Analog and RS-232 serial output with 4,000-count range. Unit operates continuously when power is supplied. |  | 2019-11-07 AWWr
WETLabs |  |  | OKÄND |  |  |  | Stored | WL2 (3) | Lager Hylla 7D | [6 empty] | 2018-08-31/MN
WETLabs | 600m |  | FLPERT-1142 |  | 2012-08-15 | Rh | Missing | [6 empty] | Rhodamine(Rh) (520/595nm); |  | 2018-04-20/MN
WETLabs WQM
WETLabs WQM | 100m |  | 244 |  | 2017-02-08 |  | Destroyed |  | Lager Hylla 1C | CODAS01/Koster |  | Kustboj | Yes | C/T sensor WQM 100m S/N 5369, Oxygen sensor model 43I S/N 43-2904, WETLabs Fluorometer FLNTUS S/N 1987 | Recovered at Koster 2020-06-30. Most likely lead to a total loss according to Sea-Bird support 2020-08-13/MN | 2020-08-12/MN
WETLabs WQM | 100m |  | 239 |  | 2020-10-19 |  | Stored |  | Lager Hylla 1C | Koster/CODAS 02 |  | Kustboj | Yes | C/T sensor WQM 100m S/N 5196, Oxygen sensor model 43I S/N 43-0549; WETLabs Fluorometer FLNTUS S/N 2712. | Back from calibration 2020-11-17. CTD sensor changed from 5436 to 5196, DO from 3268 to 0549 and eco from 2796 to 2712. Planned for Koster 2021. | 2020-12-07/MN

## Nortek
Model | Max. depth (m) | Max measuring range (m) | SMHI Nr | Order number | Head ID | Hardware ID | Firmware version | Recorder size | Latest Calibration date | Status | Box | Place | Station | Label on instrument | Cleaned | Batteries | Function test done | Programmed | Ready for deployment | Comments | Date/Signature
AQUAPRO 400kHz
AQUAPRO 400kHz | 100m | 60-90m |  | P22037-1 | AQP 3205 | AQD 5572 | 3.40 | 9 MB | 2008-11-13 | Stored |  | Lager Hylla 5A |  |  | Yes | No | No | No | No | ODAS, can not communicate 2019-04-15 | 2019-04-25/MN
AQUAPRO 400kHz | 100m | 60-90m |  | P22967-3 | AQP 4018 | AQD 5642 | 3.43 | 9 MB | 2010-01-08 | Stored |  | Lager Hylla 8B |  | ODAS01 | Yes | No | No | No | No | ODAS, Recovered Huvudskär 2020-12-10. | 2020-12-17/MN
AQUAPRO 400kHz | 100m | 60-90m |  | P22037-2 | AQP 3390 | AQD 5566 | 3.43 | 9 MB |  | Planned |  | Lager Hylla 8B |  | ODAS03 | Yes | No |  |  |  | ODAS. Planned for Huvudskär 2021 | 2018-02-27/MN
AQUAPRO 400kHz | 100m | 60-90m |  | P25496-8 | AQP 10906 | AQD 9554 | 3.43 | 9 MB | 2020-12-10 | Stored | Squirrel | Lager Hylla 1C |  |  | Yes | No | No | No | No | CODAS. Back from service 2020-12-15. New head (SN: 10906, old SN: 5596) New endbell. | 2020-12-17/MN
AQUAPRO 400kHz | 100m | 60-90m |  | P27923-4 | AQP 7187 | AQD 12339 | 3.39 | 9 MB | 2014-12-30 | Stored |  | Lager Hylla 1C |  |  | Yes |  |  |  |  | CODAS, Never used. | 2018-02-08/EU
AQUAPRO 400kHz | 100m | 60-90m |  | P25496-9 | AQP 5595 | AQD 9550 | 3.43 | 9 MB |  | Planned | Squirrel | Lager Hylla 1C |  | CODAS02 | Yes | No | Yes | Yes | No | CODAS, Recovered from Tångesund June2017.Planned for Koster 2021. | 2020-12-07/MN
AQUAPRO 600kHz (APV-instrument)
AQUAPRO 600 kHz | 200m | 30-40m |  | P014/04 | AQP 1115 | AQD 1358 | 1.11 | 9 MB |  | Stored |  | Lager Hylla 4A | [7 empty] | Recovered from Halsfjärden dec 2012. Service required? | 2018-11-27/MN
AQUAPRO 1MHz
AQUAPRO 1 MHz | 300m | 12-20m |  | P20317-2 | AQP 1826 | AQD 2115 |  | 9 MB | 2009-11-24 | Missing | [9 empty] | Found certificates but don't know where the instrument is | 2018-02-08/MN
AQUADOPP 2MHz
AQUADOPP 2MHz | 300m | - |  | P25496-1 | AQD 5600 | AQD 8370 | 3.39 | 9 MB | 2012-09-27 | Stored | Squirrel | Lager Hylla 1C |  |  | Yes | No | Yes | No | No | CODAS, Pressure sensor not installed. Recovered at Koster 2020-06-30 | 2020-08-12/MN
AQUADOPP 2Mhz | 300m | - |  | P27573 | AQD 7065 | AQD 11982 | 3.40 | 9 MB | 2014-09-29 | Deployed |  | Lager Hylla 5A |  | Hanöbukten | Yes | Yes | Yes | Yes | No | Deployed at Hanöbukten 2021-03-20 | 2021-03-20/MN
AQUADOPP 2Mhz | 300m | - |  | P26351-1 | AQD 6241 | AQD 11341 | 3.40 | 9 MB | 2013-08-01 | Deployed |  | Lager Hylla 5A | Understen | System 1/Understen | Yes | Yes | Yes | Yes | Yes | Deployed at Understen 2020-12-05 | 2020-12-07/MN
AQUADOPP 2Mhz | 300m | - |  | P25496-3 | AQD 5602 | AQD 8388 | 3.40 | 9 MB | 2012-09-28 | Planned | Squirrel | Lager Hylla 1C |  | CODAS02 | Yes | No | Yes | Yes | No | CODAS, Pressure sensor not installed. Recoverd from Tångesund June2017. Planned for Koster 2021 | 2020-12-07/MN
AQUADOPP 2Mhz | 300m | - |  | P26351-3 | AQD 6245 | AQD 11314 | 3.35 | 9 MB | 2013-08-02 | Stored |  | Lager Hylla 1C |  |  |  | No |  |  |  | Never used. Extra long, can be used with 2 battery packages. | 2018-02-08/EU
AQUADOPP 2Mhz | 300m | - |  | P26351-2 | AQD 6244 | AQD11332 | 3.40 | 9 MB | 2013-08-05 | Recovered |  | Lager Hylla 5A | Hanöbukten | System 2/Hanöbukten | Yes | Yes | Yes | Yes | Yes | Recovered at Hanöbukten 2021-03-20 | 2021-03-20/MN
AWAC 600kHz (APV-instrument) | [5 empty] | Instrument ID
AWAC 600kHz | 300m | 50m |  | P20537-1 | WAV 5346 | WPR 0464 | 1.19 AST | 82 MB | 2006-09-06 | Stored |  | Lager Hylla 8C |  |  | Yes | No | No | No | No | Recovered from Fiskeback 2018-01-11. Can not communicate | 2019-08-13/MN
AWAC 600kHz | 300m | 50m |  | P20537-4 | WAV 5345 | WPR 0476 | 1.19 AST | 82 MB | 2006-09-07 | Stored |  | Lager Hylla 8C |  |  | Yes | No | No | No | No | Recovered from Öregrund 1 2020-09-14 | 2020-09-22/RS
AWAC 600kHz | 300m | 50m |  | P21398-2 | WAV 5599 | WPR 1521 | 3.38 AST | 185 MB | 2012-10-17 | Stored |  | Lager Hylla 8C |  |  | Yes | No | No | No | No | Recovered from Fiskeback 2018-01-11 | 2018-01-12/MN
AWAC 600kHz | 300m | 50m |  | P20317 | WAV 5268 | WPR 0380 | 1.19 AST | 82 MB |  | Stored |  | Lager Hylla 8C |  |  | Yes | No | No | No | No | Recovered from Öregrund 3 2020-09-15 | 2020-09-22/RS
AWAC 600kHz | 300m | 50m |  | P21398-1 | WAV 5549 | WPR 1470 (Prolog ID: 47) | 3.38 AST (Prolog: 4.22) | 3882 MB | 2007-12-19 | Stored |  | Lager Hylla 8C |  |  | Yes | No | No | No | No | Recovered from Öregrund 2 2020-09-15 | 2020-09-22/RS
AWAC 600kHz | 300m | 50m |  | P20537-2 | WAV 5357 | WPR 0479 | 1.19 AST | 82 MB | 2006-09-06 | Stored |  | Lager Hylla 5A |  |  | Yes | No | No | No | No | Recovered from Väderöarna 2018-04-11 | 2018-11-27/MN
AWAC 600kHz | 300m | 50m | 10431 | P037/05 | WAV 5151 | WPR 0240 | 1.19 AST | 82 MB | 2012-10-16 | Stored |  | Lager Hylla 8C |  |  | Yes | No | No | No | No | Recovered Ångermanälven 2018-11-07 | 2018-08-23/MN
AWAC 600kHz | 300m | 50m |  | P20537-3 | WAV 5355 | WPR 0475 |  | 82 MB | 2006-09-07 | Missing | [9 empty] | Found certificates but don't know where the instrument is | 2018-02-08/MN
Signature 500 kHz |  |  |  |  | Head S/N | Instrument S/N
Signature 500 kHz | 300m | 60-70m | 17107 | 41302-1-634 | D-1078 | 101078 | 1.4.4093.2205_14/163 | 16 GB | 6-July-2018 | Deployed |  | Lager Hylla 5A |  |  | Yes | No | No | No | No | Deployed at P22/3892 2021-03-21 | 2021-03-21/MN
Vector |  |  |  |  | Head S/N | Instrument S/N
Vector fixed stem | 20m |  |  | P-21245 | 4546 |  |  |  |  | Stored |  | Lager Hylla 4A | [8 empty] | 2018-11-27/MN
Note on H1: Nordström Maria: Systems shipped before April 2008 should have firmware 1.XX. After April 2008 3.XX.

## YSI
Model | Max. depth (m) | Part number | Latest Calibration date | Status | Box | Place | Comments | Date/Signature
Handheld Salinity, Conductivity and Temperature System
30M/25 FT |  | 03C1279 |  | Stored |  | Lab (rum 127)

## SAIV
Model | Max. depth (m) | Serial number | Latest Calibration date | Sensors | Status | Box | Place | Additional Information | Comments | Date/Signature
SAIV Mini STD/CTD
SD204 | 500m | 433 | 2018-10-24 | C, T, DO, P, Chl, Turb | Stored |  | Lager Hylla 6A | Optional Sensor Unit (OSU106) S/N 279, Oxygen Sensor (SAIV205, OxyGuard Ocean DO Probe) S/N 263, Seapoint Chlorophyll Fluorometer (SCF3656), Seapoint Turbidity Meter. | Referensinstrument vid utsättningar. | 2020-02-07/MN
SD204 | 500m | 779 | 2018-05-30 | C, T, P, Turb | Stored |  | Lager Hylla 6A | Optional Sensor Unit (OSU106) S/N 204, Seapoint Turbidity Meter. | Referensinstrument vid utsättningar. | 2018-06-19/MN
SD204 | 500m | 578 | 2014-06-10 | C, T, P, Chl | Stored |  | Svarvrummet |  | Används ej, brukade användas på L9 | 2018-05-09/AWW
SD204 | 500m | 353 | 2015-07-10 | C, T, P | Stored |  | Svarvrummet |  | Används ej, behöver undersökas | 2018-05-09/AWW

## Aanderaa
Model | Max. depth (m) | Serial number | Latest Calibration date | Sensors | Status | Box | Place | Comments | Date/Signature
Oxygen Optode/Temp Sensor
Oxygen Optode/Temp Sensor 4175S |  | 546 |  | O, T | Stored |  |  | Sh, So or Apv? SMHI nr 10853
RCM 9 Apv-instrument
Recording Unit RCM 9 |  | 1155 |  | C, T, P | Stored |  |  | Doppler current Profiler Sensor 3920 S/N: 838, Pressure Sensor 3815A Range: 0-700kPa S/N: 801, Conductivity Cell 3619 S/N:1121, Temperature Sensor 3621 S/N: 1575
Recording Unit RCM 9 |  | 1213 |  | C, T | Stored |  |  | Doppler current Profiler Sensor 3920 S/N: 937, Conductivity Cell 3919A S/N:186, Temperature Sensor 3621 S/N: 1895
Recording Unit RCM 9 |  | 1156 |  | C, T, P | Stored |  |  | Doppler current Profiler Sensor 3920 S/N: 839, Pressure Sensor 3815A Range: 0-700kPa S/N: 796, Conductivity Cell 3619 S/N:1078, Temperature Sensor 3621 S/N: 1657

## AML
Model | Max. depth (m) | Part number | Serial number | Purchase Date | Latest Calibration date | Range | Status | Box | Place | Station | Comments | Date/Signature
Smart X
Smart X | 500m | PDC-SMX-P1S1UV-05 | 20188 | 2016-09-11 | 2016-11-08 |  | Deployed | AML | Lager Hylla 6A | Flinten 7 | Reconfigured to set up as 20226 at AML 190330. Deployed 2020-09-01 | 2020-10-09/MN
Smart X | 500m | PDC-SMX-P1S2-05 | 20226 | 2018-01-15 | 2018-02-06 |  | Stored | AML | Lager Hylla 6A |  | Recovered Flinten-7 2020-09-01 | 2020-10-09/MN
CT Xchange
CT Xchange |  | XCH-CT-RA090-n545 | 450123 | 2016-09-11 | 2018-09-11 | 0-90mS/cm, -5-+45°C | Deployed | AML | Lager Hylla 6A | Flinten 7 | Back from calibration 2018-10-17, Deployed 2020-09-01 | 2020-10-09/MN
CT Xchange |  | XCH-CT-RA090-n545 | 450172 | 2018-01-15 | 2020-11-17 | 0-90mS/cm, -5-+45°C | Stored | AML | Lager Hylla 6A |  | Recovered Flinten-7 2020-09-01. Back from calibration 2020-12-01 | 2020-12-01/MN
P Xchange
P Xchange |  | XCH-PRS-0100 | 305192 | 2016-09-11 | 2018-09-12 | 100 dBar | Deployed | AML | Lager Hylla 6A | Flinten 7 | Back from calibration 2018-10-17, Deployed 2020-09-01 | 2020-10-09/MN
P Xchange |  | XCH-PRS-0100 | 305437 | 2018-01-15 | 2020-11-13 | 100 dBar | Stored | AML | Lager Hylla 6A |  | Recovered Flinten-7 2020-09-01. Back from calibration 2020-12-01 | 2020-12-01/MN
P Xchange (ägs av SLU) |  | XCH-PRS-1000 | 306051 | 2018-12-01 | 2019-01-14 | 1000dbar | Stored |  | Svea Scientific Store | MVP |  | 2019-10-30/MN
UV Xchange
UV Xchange | 500m | XCH-UV-V | 800112 | 2016-09-11 | 2016-11-08 |  | Deployed | AML | Lager Hylla 6A | Flinten 7 | Deployed 2020-09-01 | 2020-10-09/MN
UV Xchange | 500m | XCH-UV-V | 800191 | 2018-01-15 | 2017-07-18 |  | Stored | AML | Lager Hylla 6A |  | Recovered Flinten-7 2020-09-01 | 2020-10-09/MN
LED
LED |  | SUB-G0713 | 850140 | 2016-09-11 |  | 78° 3.0 mW, -5-+55°C | Stored | AML | Lager Hylla 6A |  | Probably already used for 5000h | 2018-08-29/MN
LED |  | SUB-G0763 | 850494 | 2018-01-15 |  | 78° 1.5 mW, -5-+55°C | Stored | AML | Lager Hylla 6A |  | Recovered Flinten-7 2020-09-01: probably already used for 5000h | 2020-10-09/MN
LED |  | SUB-G0763 | 850488 | 2018-01-15 |  | 78° 1.5 mW, -5-+55°C | Deployed | AML | Lager Hylla 6A | Flinten 7 | Deployed 2020-09-01 | 2020-10-09/MN
MVP X
MVP X (ägs av SLU) | 6000m | PDC-MVPX-CTD-60 | 9042 | 2018-12-01 | 2019-01-15 |  | Stored |  | Svea Scientific Store | MVP |  | 2019-10-30/MN
Micro-T Xchange
Micro-T Xchange (ägs av SLU) |  | XCH-TMP-N545? | 404465 | 2018-12-01 | 2018-12-17 | -5-+45°C | Stored |  | Svea Scientific Store | MVP |  | 2019-10-30/MN
C Xchange
C Xchange (ägs av SLU) |  | XCH-CND-RA090 | 503812 | 2018-12-01 | 2019-01-08 | 0-90mS/cm | Stored |  | Svea Scientific Store | MVP |  | 2019-10-30/MN

## Interocean Releaser
Model | Max. depth (m) | Max. load (kg) | Weight in Air/Water | Serial number (unit) | Serial number (PCB) | Codes Release | Codes Enable | Codes Disable | Status | Box | Place | Station | Label on instrument | New electronic batteries | New release batteries | Number of releases | Days deployed | Batteries Connected/disconnected | Batteries | Battery Pack (s) | Manual | Original Ship Date | Comments | Date/Signature
111 (Acoustic Release, Yellow))
111 | 300 | 100 | 4,5kg/0,5kg | 58542288 | 6110726 | ABEH | BDF | BDFH | Deployed |  | Lager Hylla 3A | P22/3892 | Releaser GBG - 5 | 2020-05-26 | 2020-05-26 | 9 |  | Disconnected | Yes | Ext Life | 40111800000 Rev F | 2013-01-02 | Planned for P22/3892 2021-03 | 2021-03-02/MN
111 | 300 | 100 | 4,5kg/0,5kg | 59462372 | 6100641 | ABCH | BCG | BCGH | Deployed | 111 Acoustic Release (4) | Lager Hylla 3A | Understen | Releaser GBG - 2 | 2020-06-23 | 2020-06-23 | 6 | 46 | Connected | Yes | Ext Life | 40111800000 Rev G | 2013-07-02 | Deployed at Understen 2020-12-05 | 2020-12-17/MN
111 | 300 | 100 | 4,5kg/0,5kg | 59462373 | 6100675 | ABFH | BDG | BDGH | Recovered | 111 Acoustic Release (4) | Lager Hylla 3A |  | Releaser GBG - 3 | 2020-06-23 | 2017-12-05 | 5 | 270 | Disconnected | Yes | Ext Life | 40111800000 Rev G | 2013-07-02 | Recovered at Hanöbukten 2021-03-20 | 2021-03-20/MN
111 | 300 | 100 | 4,5kg/0,5kg | 59462374 | 6100674 | ACGH | CDF | CDFH | Deployed | 111 Acoustic Release (4) | Lager Hylla 3A | L9 | Releaser GBG - 1 | 2020-11-03 | 2020-11-03 | 4 |  | Connected | Yes | Ext Life | 40111800000 Rev G | 2013-07-02 | Deployed at L9 2020-12-13 | 2020-12-17/MN
111 | 300 | 100 | 4,5kg/0,5kg | 60662425 | 8240828 | ADEH | CDG | CDGH | Deployed | 111 Acoustic Release (4) | Lager Hylla 3A | Hanöbukten | Releaser GBG - 4 | 2020-09-10 | 2020-09-10 | 2 | 18 | Connected | Yes | Ext Life | 40111800000 Rev K | 2014-05-05 | Deployed at Hanöbukten 2021-03-20 | 2021-03-20/MN
1090 E (White)
1090E | 1200 | 4500 |  | 52431916 | 3200148 | ABCH | BCGH | BCG | Missing | [10 empty] | Std 2‐Pack & Ext | 40111800000 Rev F | 2010-05-17 |  | 2020-02-05/MN
1090E | 1200 | 4500 |  | 52431917 | 3200137 | ABDH | BDEH | BDE | Stored |  | Lager Hylla 3A | [8 empty] | Std 2‐Pack & Ext | 40111800000 Rev F | 2010-05-17 |  | 2020-02-05/MN
1090E | 1200 | 4500 |  | 52431918 | 3200140 | ABEH | BDFH | BDF | Stored |  | Lager Hylla 3A | [8 empty] | Std 2‐Pack & Ext | 40111800000 Rev F | 2010-05-17 |  | 2020-02-05/MN
1090E | 1200 | 4500 |  | 60662426 | 8310328 | AEFH | CEGH | CEG | Stored |  | Lager Hylla 3A | [8 empty] | Std 2‐Pack | 40111800000 Rev J | 2014-05-05 |  | 2020-02-05/MN
1100E Command & Ranging Unit |  |  |  |  | Serial number transducer (1120)
1100E |  |  |  | 09840324 | 02501332 |  |  |  | Stored |  | Lager Hylla 3A | [11 empty] | New batteries 2020-07-08 | 2020-09-22/MN
1100E |  |  |  | 08230400 | 08231384 |  |  |  | Stored |  | Lager Hylla 3A | [9 empty] | 41100010100 Rev S | 2013-12-11 | New batteries 2020-09-22 | 2020-09-28/JG
111 Tag-Line Canister |  | Length (m)
111 Tag-Line Canister |  | 150 | [6 empty] | Deployed |  | Lager Hylla 3A | L9 | [10 empty] | Deployed at L9 2020-12-13 | 2020-12-17/MN
111 Tag-Line Canister |  | 150 | [6 empty] | Deployed |  | Lager Hylla 3A | Hanöbukten | [10 empty] | Deployed at Hanöbukten 2021-03-20 | 2021-03-20/MN
111 Tag-Line Canister |  | 150 | [6 empty] | Recovered |  | Lager Hylla 3A | [11 empty] | Recovered at Hanöbukten 2021-03-20 | 2021-03-20/MN
111 Tag-Line Canister |  | 100 | [6 empty] | Deployed |  | Lager Hylla 3A | P22/3892 | [10 empty] | Deployed at P22/3892 2021-03-21 | 2021-03-21/MN
111 Tag-Line Canister |  | 300 | [6 empty] | Deployed |  | Lager Hylla 3A | Understen | [10 empty] | Deployed at Understen 2020-12-05 | 2020-12-17/MN
Note on O1: ocean.gbg11: Standard: 2st Lithium C cell 7Ah = 6månader Extended: 4st Lithium C cell 14Ah (blå) = 12månader
Note on P1: ocean.gbg11: 6st Lithium 3V CR123A (vit) =30 release/reset cycles
Note on Q1: ocean.gbg11: 6st Lithium 3V CR123A (vit) =30 release/reset cycles
Note on R1: ocean.gbg11: Standard: 2st Lithium C cell 7Ah = 6månader Extended: 4st Lithium C cell 14Ah (blå) = 12månader

## Annat
Model | Max. depth (m) | SMHI Nr | Job number | Order number | Part number | Serial number | Head ID | Hardware ID | Firmware version | Recorder size | Range | Latest Calibration date | Sensors | DOS Attached | Pump Attached | T sensor attached | C sensor attached | WETlabs sensor attached | Clamp size | Status | Box | Place | Station | Label on instrument | Attached to MicroCAT | Attached to SeaCAT | Additional Information | Comments | Date/Signature
Biospherical Instruments
QSP2300 |  | 012065 |  |  |  | 70154 | [6 empty] | Irradiance | [10 empty] | Fluorescence CTD |  |  | Spare instrument. Kontakt av typen XSG (gamla typen) | Place: Johan Kronsell desk | 200128/AWWr
QSP2350 | [5 empty] | 70374 | [6 empty] | Irradiance | [8 empty] | Svea CTD Hangar | Rosett |  |  |  | MCBH-kontakt |  | 200128/Jkro
QSP2350 | [5 empty] | 70711 | [5 empty] | 2018-10-29 | Irradiance | [13 empty] | MCBH-kontakt |  | 200128/Jkro
Teledyne Benthos
PSA-916 | 6000m |  |  |  |  | 1251 | [5 empty] | 2007-06-12 | Altimeter | [8 empty] | Svea CTD Hangar | Rosett | Altimeter |  |  |  | Back from Svea. Do not work. Send in for serivce | 2019-06-19/MN
RD Instruments
Rio Grande 600kHz (WHR600-1) |  | 009125 |  |  |  | 2449 | [13 empty] | Stored |  | Lager Hylla 5A | [6 empty] | 2018-11-27/MN
Nemo Wave Processing module (NEMO12SC-I) | [5 empty] | 007 | [13 empty] | Stored
Acoustic Doppler current Profiler (717-3001-00 REV K) | [5 empty] | 3258 | [13 empty] | Stored
Alpha Tracka II
Alpha Tracka II | [19 empty] | Stored |  | Lager Hylla 4A | [6 empty] | 2018-11-27/MN
RINKO III
ARO-CAV (ägs av SLU) | [5 empty] | 0344 | [5 empty] | 2018-07-13/2018-10-18 | T, DO | [6 empty] | Stored |  | Svea Scientific Store | MVP | [5 empty] | 2019-10-30/MN
Valeport
VA500 | [5 empty] | 71991 | [5 empty] | 2019-08-30 | Altimeter | [6 empty] | Stored |  | Svea SMHI Store |  |  |  |  | Tiil Rosett. Range 100m
Ocean Origo
Sea Turtle V5TP | [5 empty] | 5004 | [5 empty] | 2018-12-18 | P | [10 empty] | TP04
Sea Turtle V5TP | [5 empty] | 5005 | [5 empty] | 2018-12-09 | P | [10 empty] | TP05
Sonardyne
Type 7967 Command Unit | [5 empty] | 245828-010
Barometrar
Vaisala PA11A (referensbarometer) |  | 005355 |  |  |  | 463204 | [5 empty] | 2020-04-23 | [7 empty] | Stored |  | Lager Hylla 6A | [5 empty] | I samma box som Provtryckningsutrustning | 2020-04-29/MN
Vaisala PA11A |  | 002752 |  |  |  | 393288 | [5 empty] | 2017-02-06 | [7 empty] | Stored | Barometrar | Lager Hylla 6A | [5 empty] | Fungerar inte | 2020-04-30/MN
Vaisala PTB220 ABA2A2A1AB | [5 empty] | C2510006 | [5 empty] | 2020-04-23 | [7 empty] | Stored |  | Verkstad | [5 empty] | Vid Ocean.gbg11 | 2020-04-29/MN
RM Young 61302V | [5 empty] | BPA 8887 | [13 empty] | Stored | Barometrar | Lager Hylla 6A | [6 empty] | 2020-04-30/MN
Deepoid
Tilly | 100m | [18 empty] | Stored |  | Lager Hylla 6A |  |  |  |  | Tilt sensor, to be attached to bottom rig | Undervattensenhet skickad till tillverkaren för reparation 2020-09-17. Tillbaka 2020-11, ej testad | 2020-12-17/MN
Note on J1: Nordström Maria: Systems shipped before April 2008 should have firmware 1.XX. After April 2008 3.XX.

## Havsboj
Model | Part number | Serial number | Firmware version | Recorder size | Latest Calibration date | Status | Box | Place | Station | Label on instrument | Additional Information | Comments | Date/Signature
Buoy
Watchmate |  | 00120 |  |  |  | Recovered |  |  | ODAS01 | ODAS01 |  | Recovered at Huvudskär 2020-12-10 | 2020-12-17/MN
Watchmate |  | 00140 |  |  |  | Deployed |  | Förråd/Lager | ODAS03 | ODAS03 |  | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
Watchmate (Progatec, kopia) | [5 empty] | Stored |  | Förråd/Lager |  |  |  |  | 2020-04-14/MN
Water Level Sensor
WLS |  | 750 |  |  |  | Recovered |  |  | ODAS01 |  |  | Recovered at Huvudskär 2020-12-10 | 2020-12-17/MN
WLS |  | 740 |  |  |  | Deployed |  |  | ODAS03 |  |  | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
Power and distribution module
MET PDB |  | 00283 |  |  |  | Recovered |  |  | ODAS01 |  |  | Recovered at Huvudskär 2020-12-10 | 2020-12-17/MN
MET PDB |  | 00264 |  |  |  | Deployed |  |  | ODAS03 |  |  | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
Triaxys
TAS |  | 03490 |  |  |  | Stored | [6 empty] | 2020-12-17/MN
TAS |  | 03170 |  |  |  | Deployed |  |  | ODAS03 |  |  | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
TAS |  | 03500 |  |  |  | Recovered |  |  | ODAS01 |  |  | Recovered at Huvudskär 2020-12-10 | 2020-04-14/MN
TAS |  | 03180 |  |  |  | Stored |  | Lager Hylla 5B |  |  |  |  | 2020-04-14/MN
AIS Transceiver
L3 AtoN AIS |  | 000895592 |  |  |  | Recovered |  |  | ODAS01 |  | ODAS HUVUDSKAR OST, MMSI 992654001 | Recovered at Huvudskär 2020-12-10 | 2020-12-17/MN
L3 AtoN AIS | ATN01-100-00 | 000631179 |  |  |  | Deployed |  |  | ODAS03 |  | ODAS HUVUDSKAR OST, MMSI 992654001 | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
L3 AtoN AIS |  | 001028004 |  |  |  | Stored |  | Lager Hylla 5B |  |  | ODAS HUVUDSKAR OST, MMSI 992654001 |  | 2020-04-14/MN
Watchman
WMAN |  | 00249 |  |  |  | Recovered |  |  | ODAS01 |  | ODAS 01, 9605003644C095C6 | Recovered at Huvudskär 2020-12-10 | 2020-12-17/MN
WMAN |  | 00251 |  |  |  | Deployed |  |  | ODAS03 |  | ODAS 03, 810001d749611352, SMHI NR:15345 | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
WMAN |  | 00274 |  |  |  | Stored | [6 empty] | 2020-04-14/MN
WMAN |  | 00275 |  |  |  | Stored |  | Lager Hylla 5B |  |  | Inget ODAS-nummer, TEST275 |  | 2020-04-14/MN
WMAN |  | 00250 |  |  |  | Stored |  | Lager Hylla 5B |  |  | ODAS 05, 810001d74961136f |  | 2020-04-14/MN
Iridium modem
SBD Transceiver 9601 |  | 300034013615500 |  |  |  | Recovered |  |  | ODAS01 |  | WMAN 249 | Recovered at Huvudskär 2020-12-10 | 2020-12-17/MN
SBD Transceiver 9601 |  | 300034013708580 |  |  |  | Deployed |  |  | ODAS03 |  | WMAN 251 | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
SBD Transceiver 9601 |  | 300034012999670 |  |  |  | Stored |  |  |  |  | WMAN 274 |  | 2020-04-14/MN
SBE IMM modem
IMM RS-232/LLS | 90554.1 | 70000919 |  |  |  | Recovered |  |  | ODAS01 |  | WMAN 249 | Recovered at Huvudskär 2020-12-10 | 2020-12-17/MN
IMM RS-232/LLS | 90554.1 | 70000749 |  |  |  | Deployed |  |  | ODAS03 |  | WMAN 251 | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
IMM RS-232/LLS | 90554.1 | 70001006 |  |  |  | Stored |  |  |  |  | WMAN 274 |  | 2020-04-14/MN
IMM Head
ICC 100 turns 90578.082 | [5 empty] | Recovered |  |  | ODAS01 |  |  | Recovered at Huvudskär 2020-12-10 | 2020-12-17/MN
ICC 100 turns 90578.082 | [5 empty] | Deployed |  |  | ODAS03 |  | New 2020-08 | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
Radar Reflector
Navinordic/Plastimo 39099 | [5 empty] | Recovered |  |  | ODAS01 |  |  | Recovered at Huvudskär 2020-12-10 | 2020-12-17/MN
Navinordic/Plastimo 39099 | [5 empty] | Recovered |  |  | ODAS01 |  |  | Recovered at Huvudskär 2020-12-10 | 2020-12-17/MN
Navinordic/Plastimo 39099 | [5 empty] | Recovered |  |  | ODAS01 |  |  | Recovered at Huvudskär 2020-12-10 | 2020-12-17/MN
Navinordic/Plastimo 39099 | [5 empty] | Deployed |  |  | ODAS03 |  |  | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
Navinordic/Plastimo 39099 | [5 empty] | Deployed |  |  | ODAS03 |  |  | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
Navinordic/Plastimo 39099 | [5 empty] | Deployed |  |  | ODAS03 |  |  | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
Solar panel
OSDA SOLAR - ODA50-18-M (50W) |  | 6680096409550055 |  |  |  | Recovered |  |  | ODAS01 | SOLAR 1 |  | Recovered at Huvudskär 2020-12-10 | 2020-12-17/MN
OSDA SOLAR - ODA50-18-M (50W) |  | 6680096409550080 |  |  |  | Recovered |  |  | ODAS01 | SOLAR 2 |  | Recovered at Huvudskär 2020-12-10 | 2020-12-17/MN
OSDA SOLAR - ODA50-18-M (50W) |  | 6680096409550033 |  |  |  | Recovered |  |  | ODAS01 | SOLAR 3 |  | Recovered at Huvudskär 2020-12-10 | 2020-12-17/MN
TUV NORD - BL50M (50W) |  | 5301140505005 |  |  |  | Deployed |  |  | ODAS03 | SOLAR 1 |  | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
TUV NORD - BL50M (50W) |  | 5301140505002 |  |  |  | Deployed |  |  | ODAS03 | SOLAR 2 |  | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
TUV NORD - BL50M (50W) |  | 5301140505003 |  |  |  | Deployed |  |  | ODAS03 | SOLAR 3 |  | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
Barometer
Vaisala PTB 110 | 4B0AB | F4950020 |  |  |  | Stored | Barometrar | Lager Hylla 6A |  |  |  | Visa 0,4 hPa över ref baro vid test 2020-04-29/MN | 2020-04-29/MN
Vaisala PTB 110 4B1AB | 4B0AB | F1450076 |  |  | 2018-04-16 | Destroyed |  | Verkstad |  |  |  | Läckt in vatten i box på ODAS01, antagligen efter upptag på däck. Visar 715,7 när ref baro visar 1008,8, 2020-04-29/MN | 2020-04-29/MN
Vaisala PTB 110 | 4B0AB | E4740005 |  |  |  | Recovered |  |  | ODAS01 |  |  | Recovered at Huvudskär 2020-12-10 | 2020-12-17/MN
Vaisala PTB 110 | 4B1AB | S4610184 |  |  | 2020-11-09 | Deployed |  |  | ODAS03 |  | New from supplier 2020 | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
ASI-GPS antenna
Garmin GA30 | [5 empty] | Recovered |  |  | ODAS01 |  |  | Recovered at Huvudskär 2020-12-10 | 2020-12-17/MN
Garmin GA30 | [5 empty] | Destroyed | [5 empty] | Korrosionsskador vid kontakt, funktion osäker rangeras ut. | 2020-12-09/EU
Garmin GA30 | [5 empty] | Deployed |  |  | ODAS03 |  | Ny | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
GPS antenna
Garmin GPS16x-HVS |  | 1A4152744 |  |  |  | Recovered |  |  | ODAS01 |  |  | Recovered at Huvudskär 2020-12-10 | 2020-12-17/MN
Garmin GPS16x-HVS |  | 1A4152743 |  |  |  | Deployed |  |  | ODAS03 |  |  | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
Garmin GPS16x-HVS |  | 1A4152738 |  |  |  | Stored |  | Lager Hylla 5B |  |  | Ny |  | 2020-04-14/MN
Garmin GPS16x-HVS |  | 1A4152733 |  |  |  | Stored |  | Lager Hylla 5B |  |  | Ny |  | 2020-04-14/MN
Garmin GPS16x-HVS |  | 1A4152735 |  |  |  | Deployed |  | Lager Hylla 5B | OCSY TEST GBG |  | Använd |  | 2020-12-09/EU
AIS-VHF antenna
COMROD AV51P4 (151-162 MHz) |  | May 16 |  |  |  | Recovered |  |  | ODAS01 |  |  | Recovered at Huvudskär 2020-12-10 | 2020-12-17/MN
COMROD AV51P4 (151-162 MHz) |  | Jul 16 |  |  |  | Deployed |  |  | ODAS03 |  |  | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
Iridium antenna
Aero Antenna Technology |  | 37816 |  |  |  | Recovered |  |  | ODAS01 |  |  | Recovered at Huvudskär 2020-12-10 | 2020-12-17/MN
Aero Antenna Technology |  | 32424 |  |  |  | Deployed |  |  | ODAS03 |  |  | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
Aero Antenna Technology |  | 42555 |  |  |  | Stored |  | Lager Hylla 5B |  |  | Använd |  | 2020-04-14/MN
Aero Antenna Technology |  | 31729 |  |  |  | Stored |  | Lager Hylla 5B |  |  | Använd, svårt se S/N på ettiketten, så inte säkert det stämmer |  | 2020-04-14/MN
Air temperature sensor
RM Young YSI |  | 270 |  |  |  | Recovered |  |  | ODAS01 |  |  | Recovered at Huvudskär 2020-12-10 | 2020-12-17/MN
RM Young YSI |  | 271 |  |  |  | Deployed |  |  | ODAS03 |  |  | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
RM Young YSI |  | 272 |  |  |  | Stored |  | Lager Hylla 5B |  |  | Använd |  | 2020-04-14/MN
Buoy light
Sealite Pty. Ltd SL-70-Y |  | 532953 |  |  |  | Recovered |  |  | ODAS01 |  | FL(5)20s, B:F, A:9, 2:ON, 1:ON | Recovered at Huvudskär 2020-12-10 | 2020-12-17/MN
Sealite Pty. Ltd SL-70-Y |  | 421596 |  |  |  | Deployed |  |  | ODAS03 |  | FL(5)20s, B:F, A:9, 2:ON, 1:ON | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
Sealite Pty. Ltd SL-70-Y |  | 302320 |  |  |  | Deployed |  |  | ODAS03 |  | FL(Y)3s, B:0, A:A, 2:ON, 1:OFF | Deployed at Huvudskär 2021-03-17 | 2021-03-20/MN
Sealite Pty. Ltd SL-70-Y |  | 555692 |  |  |  | Stored |  | Lager Hylla 5B |  |  | Ny |  | 2020-04-14/MN
Wind sensor
RM Young 05103 |  | WM126106 |  |  |  | Deployed |  |  | ODAS03 |  |  | Deployed at Huvudskär 2021-03-17 | 2020-12-17/MN
RM Young 05106 |  | WM173292 |  |  | 2019-10-31 | Stored |  |  |  |  | Ny, inköpt 2020-04-03, kablage fixat och testat |  | 2020-04-14/MN
RM Young 05106 | [5 empty] | Stored |  |  |  |  | Ny, inköpt 2020-12, kablage ej fixat och testat |  | 2021-02-18/MN
PDB/Sunsaver Interface
2225 PDB/Sunsaver Interface |  | 00440 |  |  |  | Stored |  | Lager Hylla 5B |  |  | Ny |  | 2020-04-14/MN
Note on D1: Nordström Maria: NORTEK: Systems shipped before April 2008 should have firmware 1.XX. After April 2008 3.XX.

## Kustboj
Model | Part number | Serial number | Firmware version | Recorder size | Latest Calibration date | Status | Box | Place | Station | Label on instrument | Additional Information | Comments | Date/Signature
Buoy
Mobilis DB500 | [5 empty] | Stored |  |  | CODAS01 | CODAS01 |  | Recovered at Koster 2020-06-30 | 2020-08-13/MN
Mobilis DB500 | [5 empty] | Planned |  |  | CODAS02/Koster | CODAS02 |  |  | 2020-12-17/MN
TMBB
TMBB20 |  | 1115 |  |  |  | Planned |  |  | CODAS02/Koster |  | Sim card nr: 89460 11013 09060 61108, sim phone number: 07190000454822 | Recovered at Koster 2020-06-30 | 2020-12-17/MN
TMBB20 |  | 1118 |  |  |  | Stored |  |  |  |  | Sim card nr: 89460 11013 09060 30533, sim phone number: |  | 2020-04-21/MN
Solar charge controller
Steca Electronic PR1010 |  | 715818AJ005767200061 |  |  |  | Planned |  |  | CODAS02/Koster |  | Inside TMBB1115 | Recovered at Koster 2020-06-30 | 2020-12-17/MN
Steca Electronic PR1010 | [5 empty] | Stored |  |  |  |  | Inside TMBB1118 | 3412 | 2020-04-21/MN
Solar panels
Kyocera KD50SE-1P |  | 115EGG0842 |  |  |  | Stored |  |  | CODAS01 | SOLAR 1 |  | Recovered at Koster 2020-06-30 | 2020-08-13/MN
Kyocera KD50SE-1P |  | 115EGG0843 |  |  |  | Stored |  |  | CODAS01 | SOLAR 2 |  | Recovered at Koster 2020-06-30 | 2020-08-13/MN
Kyocera KD50SE-1P |  | 115EGG0838 |  |  |  | Planned |  |  | CODAS02/Koster | SOLAR 1 |  |  | 2020-12-17/MN
Kyocera KD50SE-1P |  | 115EGG0839 |  |  |  | Planned |  |  | CODAS02/Koster | SOLAR 2 |  |  | 2020-12-17/MN
Meteorology Sensor
Airmar Weather station 200WX |  | 2890616 |  |  |  | Stored |  |  | CODAS01 |  |  | Recovered at Koster 2020-06-30 | 2020-08-13/MN
Airmar Weather station 200WX |  | 3098092 |  |  |  | Planned |  |  | CODAS02/Koster |  |  |  | 2020-12-17/MN
Airmar Weather station 200WX |  | 3042731 |  |  |  | Stored |  | Lager Hylla 1B |  |  |  | Extra | 2020-06-30/MN
GSM antenna
COMROD AC21P1 |  | May -16 |  |  |  | Planned |  |  | CODAS02/Koster |  |  | Recovered at Koster 2020-06-30 | 2020-12-17/MN
COMROD AC21P1 | [5 empty] | Stored |  | Lager Hylla 1B |  |  |  | For CODAS02 | 2020-04-21/MN
Buoy light
Sealite Pty. Ltd SL-70-Y |  | 555806 |  |  |  | Stored |  |  | CODAS01 |  | FL(5)20s, B:F, A:9, 2:ON, 1:ON, Battery 2016-36 | Recovered at Koster 2020-06-30 | 2020-08-13/MN
Sealite Pty. Ltd SL-70-Y |  | 3098092 |  |  |  | Planned |  |  | CODAS02/Koster |  | FL(5)20s, B:F, A:9, 2:ON, 1:ON |  | 2020-12-17/MN
Note on D1: Nordström Maria: NORTEK: Systems shipped before April 2008 should have firmware 1.XX. After April 2008 3.XX.

## Listor
Sensors | Status | Place | Box | Antifouling | Cleaned | Data Download | Batteries | Diagnostics done | Programmed | Ready for deployment | Wiper | Clamp size | Batteries Connected/disconnected
Altimeter | Stored | Förråd/Lager | Adela (4) | Yes | Yes | Yes | Yes | Yes | Yes | Yes | Yes | No clamps | Connected
C ,T, DO, P, Chl, Turb, PC, Irradiance, Depth | Planned | Lab (rum 127) | Aldebaran (5) | No | No | No | No | No | No | No | No | 8mm | Disconnected
C, T | Deployed | Lager Hylla 1A | Alpha Centauri (5) | Copper |  |  | Old |  |  |  |  | 10mm
C, T, DO | Recovered | Lager Hylla 1B | AML | Plastic | [7 empty] | 12mm
C, T, DO, Chl, Turb | Destroyed | Lager Hylla 1C | Bellatrix (5) | [8 empty] | 16mm
C, T, DO, Chl, Turb, PC | Lost | Lager Hylla 1D | Birk (5) | [8 empty] | 3/8"
C, T, DO, P | Missing | Lager Hylla 2A | Borka (5) | [8 empty] | 7/16"
C, T, DO, P, Chl, Turb | Loaned | Lager Hylla 2B | Deneb (5)
C, T, P | Calibration | Lager Hylla 2C | Furud (5)
C, T, P, Chl |  | Lager Hylla 2D | Lovis (2)
C, T, P, Turb |  | Lager Hylla 3A | Mattis (4)
Chl |  | Lager Hylla 3B | Polaris (4)
Chl, Turb |  | Lager Hylla 3C | P-Rigel (4)
Irradiance |  | Lager Hylla 3D | Proxima (2)
O, T |  | Lager Hylla 4A | Rigel (5)
P |  | Lager Hylla 4B | Ronja (4)
PC |  | Lager Hylla 4C | Sirius (5)
Rh |  | Lager Hylla 4D | Squirrel (4)
T, DO |  | Lager Hylla 5A | Trappist (4)
|  | Lager Hylla 5B | Undis (5)
|  | Lager Hylla 5C | Vega (2)
|  | Lager Hylla 5D | WL1 (3)
|  | Lager Hylla 6A | WL2 (3)
|  | Lager Hylla 6B | WL3 (3)
|  | Lager Hylla 6C | 111 Acoustic Release (4)
|  | Lager Hylla 6D | Barometrar
Lager Hylla 7A
Lager Hylla 7B
Lager Hylla 7C
Lager Hylla 7D
Lager Hylla 8A
Lager Hylla 8B
Lager Hylla 8C
Lager Hylla 8D
Lager Hylla 9A
Lager Hylla 9B
Lager Hylla 9C
Lager Hylla 9D
Lysekil Hamnmagasin
Svarvrummet
Svea CTD Hangar
Svea Scientific Store
Svea Sea Water lab
Svea SMHI Store
Tavastland
Verkstad

## Sensor_info
cnv-namn | cnv-code | Instrument ID | Instrumenttillverkare | Instrumentmodell | Sensortillverkare | Sensormodell | Sensor serienummer | Parameter enkel | Parameter (Kod) | Enhet | Beräknad parameter (ja/"tom") | Referens till beräkningsmetod | Rapporteringsgräns (fritext) | Mätområde (fritext) | Mätosäkerhet (fritext) | Mätfrekvens (Hz) | Mätmetod kommentar
CNV_NAME | CNV_CODE | INSTRUMENT_ID_PREFIX | INSTRUMENT_PROD | INSTRUMENT_MOD | SENSOR_PROD | SENSOR_MOD | SENSOR_ID | PARAM_SIMPLE | PARAM | MUNIT | CALCULATED | CALCULATE_REF | LMQNT | RANA | UNCERT | FREQ | MET_COMNT
Depth [true depth, m] | depFM | [6 empty] | Djup | DEPH | m | Yes | Calculated from true density profile |  |  |  |  | calculated from density and pressure
Pressure | prDM | SBE | Seabird | 911plus | Paroscientific | Digiquartz 8000-series | 1044, 1387, 0745, 1464 | Tryck | PRES_CTD | dbar |  |  | 0 | 0-1400 m | 0.01% Typical Accuracy | 24
Pressure | prdM | SBE | Seabird | 19plus | Seabird | SeaCAT | 6164, 6929, 8438 | Tryck | PRES_CTD | dbar |  |  |  | 0-350 | +/- 0.1% of full scale range | 4 | Typical stability: Strain-gauge ± 0.1% of full scale range per year. Resolution: Strain-gauge 0.002% of full scale range;
Conductivity | c0S/m, c1S/m |  |  |  | Seabird | SBE 4C | 3941, 3943, 2997, 2994, 4829, 4824, 5060 | Konduktivitet | CNDC_CTD | S/m |  |  | 0 | 0.0-7.0 S/m | ± 0.0003 S/m | 24
Conductivity | c0S/m, c1S/m |  |  |  | Seabird | SeaCAT | 6164, 6929, 8438 | Konduktivitet | CNDC_CTD | S/m |  |  |  | 0 to 9 S/m | ± 0.0005 S/m | 4 | Typical stability: 0.0002 °C per month. Resolution: 0.0001 °C
Salinity | sal00, sal11 | [6 empty] | Salthalt | SALT_CTD | psu | Yes | UNESCO 1983 (EOS 80) | 2 |  |  |  | calculated from conductivity, temperature and pressure
Temperature | t090C, t190C |  |  |  | Seabird | SBE 3plus | 5351, 5355, 6378, 6384, 4361, 4387, 6596 | Temperatur | TEMP_CTD | °C (ITS-90) |  |  | -5.0 | -5.0 - +35 °C | ± 0.001 °C | 24
Temperature | tv290C |  |  |  | Seabird | SeaCAT | 6164, 6929, 8438 | Temperatur | TEMP_CTD | °C (ITS-90) |  |  |  | -5 to +35 T5 | ± 0.01 °C | 4 | Typical Stability: 0.0003 S/m per month. Resolution: 0.00005 S/m typical
Oxygen | sbeox0ML/L, sbeox1ML/L |  |  |  | Seabird | SBE 43 | 0732, 1243, 0732, 2126, 2127, 3823, 3829, 4152, 4649 | Syrgashalt | DOXY_CTD | ml/l |  |  | 0.1 | 120% of surface saturation in all natural waters, fresh and salt |  | 24* | Based on signal 0-5VDC. *Note that a SBE43 mounted on a SBE19 has a sampling frequency of 4Hz.
Oxygen, SBE 43 [% saturation] | sbeox0PS, sbeox1PS | [6 empty] | Syremättnad | DOXY_SAT_CTD | % | Yes
PAR/Irradiance, Biospherical/Licor | par |  |  |  | Biospherical | QSP-2000, QSP-2350 | 70374, 70711, 70397, 2140, 2112, 70785 | PAR | PAR_CTD | µE/(m2 sec) |  |  | 0.14 | 0.14 µE/(m2 ·sec) to 5000 µE/(m2 ·sec)
Fluor | flECO-AFL, flECO-AFL1 |  |  |  | WET Labs | ECO-FLNTU-RTD | FLNTURTD-2288, FLNTURT-115, FLNTURT-1899, FLNTURT-2470, FLNTURT-9297 | Klorofyllfluorescens* | CHLFLUO_CTD* | mg/m3 |  |  | 0 | 0-50 µg/I | 0.025 µg/I | 8
Turbidity | turbWETntu0 |  |  |  | WET Labs | ECO-FLNTU-RTD | FLNTURTD-2288, FLNTURT-115, FLNTURT-1899, FLNTURT-2470, FLNTURT-9297 | Turbiditet | TURB_CTD | NTU |  |  | 0 | 0-25 NTU | 0.013 NTU | 8
Fluor | flECO-AFL, flECO-AFL1 |  |  |  | WET Labs | ECO-FLPC-RTD | FLPCRTD-2299 | Phycocyaninfluorescens* | PHYC_CTD* | ppb |  |  | 0 | 0-230 ppb | 0.03 ppb | 8
Altimeter | altM |  |  |  | Teledyne | PSA-916 | 53447, 1251 | Altimeter | ALTM_CTD | m |  |  | 1 | 1-100 m | 2.5 cm
Altimeter | altM |  |  |  | Valeport | VA500 | 87405, 71991 | Altimeter | ALTM_CTD | m |  |  | 1 | 1-100 m | 2.5 cm
Descent Rate | dz/dtM | [6 empty] | fallhastighet | DESC_RATE_CTD | m/s | Yes | derivative of pressure with respect to time |  |  |  |  | calculated in seabird dataprocessing by using the derivative of pressure with respect to time
Number of scans per bin | nbin | [6 empty] | scans per bin | SCAN_BIN_CTD
Scan Count | scan | [6 empty] | scan nummer | SCAN_CTD
Pump Status | pumps |  |  |  | Seabird | SBE 5T |  | pump status | PUMP_CTD
Density | sigma-é00, sigma-é11 | [6 empty] | Densitet | SIGMA_THETA_CTD | kg/m3 | Yes | UNESCO 1983 (EOS 80) |  |  |  |  | calculated from salinity, temperature and pressure
Density | density00, density11 | [6 empty] | Densitet | DENS_CTD | kg/m3 | Yes | UNESCO 1983 (EOS 80) |  |  |  |  | calculated from salinity, temperature and pressure
Sound Velocity | svCM, svCM1 | [6 empty] | Ljudhastighet | SVEL_CTD | m/s | Yes | UNESCO 1983 (EOS 80) Chen-Millero |  |  |  |  | calculated from salinity, temperature and pressure
Salinity, Practical, Difference, 2 - 1 [PSU] | secS-priS | [6 empty] | Salthalt diff | SALT_DIFF_CTD | psu | Yes
Temperature Difference, 2 - 1 [ITS-90, deg C] | T2-T190C | [6 empty] | Temperatur diff | TEMP_DIFF_CTD | °C (ITS-90) | Yes
Time, NMEA | timeQ | [6 empty] | Scan tid (sekunder sedan 2000-01-01) | SCAN_TIME_CTD | seconds | Yes
Time, Elapsed | timeS | [6 empty] | Förfluten tid (sekunder sedan system upload time) | SCAN_TIME_ELAP_CTD | seconds | Yes